## Gap Analysis Worksheet
Gap Analysis Worksheet: Harmonizing FDA QSR, QMSR, and ISO 13485:2016
Learn more about The FDA Group's QMSR support services » Purchase ISO 13485:2016 » Purchase ISO 14971:2019 »
Item | Assessment Detail | ISO 13485:2016 | 21 CFR Part 820 FDA QSR | 21 CFR Part 820 FDA QMSR | Auditor Actions (to comply with QMSR/ISO) | Observation | Evidence | NC, OFI, PP, or A?
Management Controls
1 | Verify that the Quality Manual comprehensively outlines the scope of the QMS, including justified exclusions, confirms that all necessary QMS procedures are documented or referenced, and clearly describes the interaction between QMS processes. | ISO 13485:2016: 4.1, 4.2.2 |  | ISO 13485:2016: 4.1, 4.2.2 | - Review Quality Manual: Ensure it comprehensively outlines the scope of the QMS, including any justified exclusions. Verify that it documents or references all necessary QMS procedures. - Scope and Justifications: Check that the Quality Manual includes detailed information on the scope of the QMS, and provides justifications for any exclusions or non-applications of the standard. - Process Interactions: Confirm that the Quality Manual clearly describes the interaction between various processes within the QMS. - Documented Procedures: Verify that the organization has established, implemented, and maintains the required documented procedures for the QMS, ensuring they are up-to-date and aligned with regulatory requirements. - Roles and Responsibilities: Check that the roles and responsibilities under the applicable regulatory requirements are documented in the Quality Manual. - Regular Review: Verify that the Quality Manual and associated procedures are subject to regular review and updates to reflect any changes in processes or regulatory requirements. - Availability and Accessibility: Confirm that the Quality Manual and documented procedures are readily available and accessible to relevant personnel.
2 | Verify criteria and methods are in place to monitor and control processes for effectiveness | ISO 13485:2016: 4.1.3(a), 4.2.1(d), 8.4 |  | ISO 13485:2016: 4.1.3(a), 4.2.1(d), 8.4 | - Document Review: Verify that all relevant documents, including procedures, records, and analysis reports, are available and up to date. Interviews: Conduct interviews with personnel responsible for monitoring and controlling processes to ensure they understand and follow the documented criteria and methods. - Process Observation: Observe key processes in action to confirm that the criteria and methods are being effectively applied. - Records Check: Review records of process monitoring, measurements, and data analysis to ensure they are accurate and demonstrate control and effectiveness.
4 | Confirm management reviews examine suitability and effectiveness of quality systems, improvements needed because of customer requirements, and resource needs | ISO 13485:2016: 4.1.3(c), 5.6.1, 5.6.3, 6.1, 8.4 | 21 CFR 820.20(c) | ISO 13485:2016: 4.1.3(c), 5.6.1, 5.6.3, 6.1, 8.4 | - Management Review Process: Ensure the management review process is comprehensively documented and includes a schedule for regular intervals (at least annually). Verify that the process assesses the suitability, adequacy, and effectiveness of the QMS. - Review Inputs: Confirm that management reviews include inputs such as customer feedback, complaints, audit results, process performance, product conformity, and corrective and preventive actions. Check for new ISO 13485:2016 and QMSR-specific inputs, including risk management and data analysis results. - Review Outputs: Verify that management review results in actionable decisions and plans for improvements, focusing on enhancing the QMS, product quality, and addressing customer requirements. Ensure documentation of decisions related to resource needs and allocation, including personnel, infrastructure, and training, considering new QMSR standards. - Resource Needs: Confirm that the review process assesses the adequacy of resources and identifies additional needs required to maintain and improve the QMS. - Data Analysis: Ensure data analysis methods are used during management reviews to evaluate process performance, product conformity, and identify improvement opportunities, aligning with ISO 13485:2016 section 8.4. - Follow-Up Actions: Verify that follow-up actions from management reviews are documented, tracked to completion, and evaluated for effectiveness, ensuring compliance with both ISO 13485:2016 requirements. - Documentation and Records: Ensure all records of management reviews, including inputs, outputs, and follow-up actions, are maintained accurately and are readily accessible for inspection. Confirm that records demonstrate compliance with the documented management review process and highlight gaps with ISO 13485:2016 and QMSR requirements. - Regular Updates: Verify that the management review process and related documentation are regularly reviewed and updated to reflect changes in regulatory requirements and organizational processes.
5 | Ensure management review addresses audit results, customer feedback, process performance, CAPAs, previous management reviews, changes to QMS, recommendations for improvement, and new or revised regulatory requirements | ISO 13485:2016: 5.6.2 |  | ISO 13485:2016: 5.6.2 | Management Review Content: - Document Request and Review: Request and review records of management reviews to ensure they are comprehensive and cover all required inputs. - Audit Results: Verify that management reviews address internal and external audit results. Ensure that the reviews consider the findings and recommendations from audits and integrate these into the QMS improvement actions. - Customer Feedback: Confirm that customer feedback, including complaints and satisfaction data, is reviewed and analyzed during management reviews. Ensure that actions are taken to address customer concerns and enhance customer satisfaction. - Process Performance: Ensure that the review includes an assessment of process performance and product conformity. Verify that key performance indicators (KPIs) and metrics are reviewed to identify trends and areas needing improvement. - Corrective and Preventive Actions (CAPAs): Verify that management reviews address the status and effectiveness of CAPAs. Ensure that the reviews include an analysis of root causes and the implementation of corrective actions to prevent recurrence. - Previous Management Reviews: Confirm that the review process considers the outputs and follow-up actions from previous management reviews. Ensure that actions identified in prior reviews have been completed and evaluated for effectiveness. - Changes to QMS: Ensure that changes to the QMS are reviewed for their impact on quality and compliance. Verify that the review includes an evaluation of changes in organizational structure, processes, and regulatory requirements. - Recommendations for Improvement: Verify that the review includes recommendations for improving the QMS and its processes. Ensure that these recommendations are based on data analysis and risk assessments. - Regulatory Requirements: Confirm that the management review considers new or revised regulatory requirements. Ensure that the organization has identified relevant changes and integrated them into the QMS to maintain compliance. Management Review Outputs: - Action Plans: Ensure that management review outputs include actionable plans for addressing identified issues and opportunities for improvement. Verify that these plans are documented and tracked to completion. - Resource Allocation: Confirm that the review results in decisions related to resource allocation, including personnel, training, infrastructure, and equipment necessary to maintain and improve the QMS. - Documentation and Records: Verify that all inputs, discussions, decisions, and actions from the management reviews are documented accurately and maintained as part of the QMS records. Ensure that these records are readily accessible for inspection and demonstrate compliance with ISO 13485:2016 requirements.
6 | Verify firm has established a Quality Manual and Quality System Procedures and Instructions that are appropriate | ISO 13485:2016: 4.1.2(a), 4.2.1(b), (c) | 21 CFR 820.5, 820.20(c), (d), (e), 820.22 | ISO 13485:2016: 4.1.2(a), 4.2.1(b), (c) | - Document Request and Review: Request the Quality Manual and Quality System Procedures and Instructions in advance. Review these documents to ensure they are comprehensive and align with regulatory requirements. - Scope and Justification: Verify that the Quality Manual includes a detailed scope of the QM) and justifications for any exclusions or non-applications of the standard. Ensure that it reflects the nature and complexity of the organization's activities. - Documented Procedures: Confirm that all necessary QMS procedures are documented or referenced within the Quality Manual. Check that these procedures cover all aspects of the QMS as required by ISO 13485:2016 and 21 CFR Part 820. - Process Interactions: Ensure the Quality Manual clearly describes the interaction between various processes within the QMS, demonstrating how each process contributes to overall quality objectives. - Regulatory Alignment: Verify that the Quality Manual and procedures are aligned with both ISO 13485:2016 and the FDA QSR (21 CFR 820). This includes adherence to specific requirements such as document controls (820.40), purchasing controls (820.50), and process controls (820.70). - Effectiveness and Maintenance: Confirm that the organization maintains the effectiveness of the documented QMS procedures. Check that there are mechanisms in place for regular review, updating, and maintaining these documents to reflect changes in regulatory requirements and organizational processes. - Roles and Responsibilities: Ensure that roles and responsibilities are clearly documented within the Quality Manual, including those for quality assurance, management review, and corrective and preventive actions. - Quality Planning: Verify that quality planning is documented, outlining the quality practices, resources, and activities relevant to the design and manufacture of devices. Ensure the plan aligns with the requirements of ISO 13485:2016 and 21 CFR Part 820. - Internal Audits: Confirm that the organization has established procedures for conducting internal quality audits. Verify that these audits are performed at defined intervals to ensure compliance with established - QMS procedures and regulatory requirements. Ensure audit results are reviewed by management and used to improve the QMS. Accessibility and Availability: Check that the Quality Manual and QMS procedures are readily accessible to relevant personnel. Ensure that these documents are available at all locations where they are needed for the implementation and maintenance of the QMS.
3 | Verify firm has established and conducts Management Reviews, at least annually | ISO 13485:2016: 5.1(d), 5.6 | 21 CFR 820.5, 820.20(c) | ISO 13485:2016: 5.1(d), 5.6 | - Management Responsibility: Verify that top management has established a policy for the quality management system (QMS) and ensures its implementation and maintenance across all levels of the organization. Check that management reviews are planned and conducted at least annually to ensure the continuing suitability, adequacy, and effectiveness of the QMS. - Management Review Process: Confirm that the organization has a documented procedure for conducting management reviews in accordance with ISO 13485:2016 section 5.6 and 21 CFR 820.20(c). Ensure that these reviews are conducted at planned intervals and include the necessary inputs and outputs as specified in the standard. - Review Inputs: Verify that the management reviews include, at a minimum, inputs such as audit results, customer feedback, process performance and product conformity data, status of preventive and corrective actions, follow-up actions from previous management reviews, changes that could affect the QMS, and recommendations for improvement. - Review Outputs: Check that the outputs of management reviews include decisions and actions related to improvement of the QMS, improvement of product related to customer requirements, and resource needs. - Documented Evidence: Ensure that the results of management reviews are documented and maintained as required by ISO 13485:2016 section 4.2.5 and 21 CFR 820.5. This documentation should include the dates of reviews, participants, review inputs, and resulting actions or decisions. - Follow-Up Actions: Confirm that there is a process for following up on actions identified during the management reviews. Verify that these actions are tracked to completion and their effectiveness is evaluated. - Frequency and Timeliness: Verify that management reviews are conducted at least annually and more frequently if necessary based on the needs and performance of the QMS.
7 | Verify firm has established Quality Plan | ISO 13485:2016: 4.2.1(d), 5.4 | 21 CFR 820.20(d) | ISO 13485:2016: 4.2.1(d), 5.4 | - Document Request and Review: Request the Quality Plan in advance.
8 | Confirm that Quality Planning addresses QMS needs and Quality Objectives | ISO 13485:2016: 5.4.2 | 21 CFR 820.20(a), (d) | ISO 13485:2016: 5.4.2 | - Quality Objectives: Verify that the Quality Plan includes clearly defined quality objectives that are consistent with the organization’s quality policy. Ensure these objectives are measurable and align with regulatory requirements. - QMS Needs: Confirm that the Quality Plan identifies and addresses the specific needs of the QMS, including processes, resources, and controls necessary to achieve the quality objectives. Ensure it considers the organizational context and the nature of the medical devices being produced.
9 | Verify firm has implemented Quality Policy and Quality Objectives | ISO 13485:2016: 4.2.1(a), 5.1(b), (c), 5.3, 5.4.1 | 21 CFR 820.20(a), (d) | ISO 13485:2016: 4.2.1(a), 5.1(b), (c), 5.3, 5.4.1 | - Employee Interviews: Interview employees to assess their understanding of the Quality Policy and how it impacts their daily activities. Confirm that they can articulate the policy and its relevance to their work. - Training and Competency: Review training records to ensure that employees have received appropriate training on the Quality Policy and Quality Objectives. Verify that training is effective and that employees are competent in their roles.
10 | Verify firm has established Quality Audit procedures and conducts audits | ISO 13485:2016: 4.2, 8.2.4 | 21 CFR 820.20(c), 820.22 | ISO 13485:2016: 4.2, 8.2.4 | - Document Request and Review: Request the Quality Audit procedures in advance and review them to ensure they comply with ISO 13485:2016 and 21 CFR Part 820 requirements. - Audit Schedule: Verify that the organization has an established audit schedule that covers all processes and areas of the QMS within a defined timeframe. Ensure that the schedule is followed and updated as necessary. - Audit Planning: Confirm that the audit procedures include detailed planning activities, such as defining the scope, objectives, and criteria for each audit. Ensure that audits are planned based on the status and importance of the processes and areas to be audited. - Auditor Training and Competency: Review training records to ensure that internal auditors have received appropriate training and are competent to conduct audits. Confirm that training includes knowledge of the audit process, QMS requirements, and relevant regulatory standards. - Documentation and Records: Ensure that all audit-related activities, including planning, execution, findings, corrective actions, and follow-ups, are documented and maintained as part of the QMS records.
11 | Ensure quality audits examine compliance and effectiveness | ISO 13485:2016: 4.1.3(c), 4.2.1(d), 8.2.4 | 21 CFR 820.22 | ISO 13485:2016: 4.1.3(c), 4.2.1(d), 8.2.4 | - Document Request and Review: Review the audit procedure to confirm that the procedure outlines the scope, frequency, and methodology for conducting audits. - Audit Criteria: Ensure that the audit criteria are clearly defined in the procedure, including references to ISO 13485:2016 sections 4.1.3(c) (ensuring the effectiveness of the QMS), 4.2.1(d) (quality planning), and 8.2.4 (internal audits), and other applicable regulatory requirements. Confirm that the criteria cover all aspects of the QMS. - Audit Records Review: Review past audit records to confirm that audits have been conducted as per the defined schedule and procedure. Verify that the records include comprehensive documentation of the audit process, findings, and conclusions.
12 | Verify that auditors are trained | ISO 13485:2016: 6.2, 8.2.4 | 21 CFR 820.22 | ISO 13485:2016: 6.2, 8.2.4 | - Training Procedures: Verify that the organization has established and maintains procedures for identifying training needs and providing training to auditors. Ensure that the procedures are documented and align with ISO 13485:2016 section 6.2 requirements for competence, awareness, and training.
13 | Ensure that audits are conducted by objective parties | ISO 13485:2016: 8.2.4 | 21 CFR 820.22 | ISO 13485:2016: 8.2.4 | - Independence of Auditors: Review past audit records to verify that audits were conducted by objective parties. Ensure that the records clearly document the auditors involved and their lack of direct responsibility for the audited areas.
14 | Confirm quality audits are linked to CAPA | ISO 13485:2016: 8.2.4 | 21 CFR 820.22, 820.100 | ISO 13485:2016: 8.2.4 | - CAPA Procedure Compliance: Review past audit records to confirm that nonconformities and areas for improvement identified during audits have been appropriately documented and linked to CAPA actions. Verify that these records include detailed information on the findings and the corresponding CAPA actions taken.
15 | Review organizational structure of firm; confirm resources are available to support processes | ISO 13485:2016: 4.1.3(b), 5.1(e), 5.5.1, 5.5.2, 6.1, 6.2 | 21 CFR 820.20(b), 820.25 | ISO 13485:2016: 4.1.3(b), 5.1(e), 5.5.1, 5.5.2, 6.1, 6.2 | - Organizational Structure: Verify that the organizational structure is clearly defined and documented. Confirm that it includes all necessary roles and responsibilities to support the QMS and that these roles are appropriately staffed. - Resource Availability: Confirm that the organization has allocated sufficient resources to support the implementation, maintenance, and improvement of the QMS. Ensure that resource planning includes considerations for personnel, infrastructure, and work environment as required by ISO 13485:2016 section 6.1.
16 | Verify firm has defined a management representative with executive responsibility for implementing and reporting quality management system | ISO 13485:2016: 5.1, 5.5.1, 5.5.2, 6.1, 6.2 | 21 CFR 820.20(b)(3), 820.25 | ISO 13485:2016: 5.1, 5.5.1, 5.5.2, 6.1, 6.2 | - Appointment Documentation: Request and review documentation that identifies the appointed management representative. Verify that the management representative has been formally appointed with documented authority and responsibility for ensuring the QMS is implemented and maintained as required by ISO 13485:2016 sections 5.5.2 and 21 CFR 820.20(b)(3). - Executive Responsibility: Confirm that the management representative has executive responsibility and is empowered to make decisions regarding the QMS. Ensure that this role includes the authority to implement changes, oversee quality management activities, and report on the performance of the QMS to top management as required by ISO 13485:2016 sections 5.1 and 5.5.1. Ask the management representative to identify their specific responsibilities, including: --» Changes to Procedures, Device Designs, and Manufacturing Processes: Verify that the management representative oversees and approves changes to QMS procedures, device designs, and manufacturing processes to ensure they meet regulatory requirements as required by ISO 13485:2016 sections 5.5.1 and 6.1. --» Review of Quality Audit Results: Confirm that the management representative reviews quality audit results and ensures that audit findings are addressed through appropriate corrective actions as required by ISO 13485:2016 section 8.2.4 and 21 CFR 820.22. --» Oversight and Interaction with CAPA Activities: Ensure that the management representative oversees and interacts with CAPA activities, verifying that nonconformities are identified, investigated, and resolved effectively as required by ISO 13485:2016 sections 8.2.4 and 8.5.2 and 21 CFR 820.100.
17 | Verify appropriate responsibilities, authority, and resources are in place for quality system activities | ISO 13485:2016: 5.1(e), 5.5.1, 5.5.2, 6.1, 6.2 | 21 CFR 820.5(b)(1)-(2), 820.20(b), 820.25 | ISO 13485:2016: 5.1(e), 5.5.1, 5.5.2, 6.1, 6.2 | - Interview Management Representative: Interview the management representative to discuss the allocation of resources and the assignment of responsibilities and authority within the organization. Confirm that they have sufficient authority to oversee and implement the QMS as required by ISO 13485:2016 sections 5.1(e) and 5.5.1 - Resource Allocation: Confirm that the organization has allocated adequate resources, including personnel, equipment, and infrastructure, to support the implementation and maintenance of the QMS. Verify that resource planning includes considerations for competence, training, and awareness of personnel as required by ISO 13485:2016 sections 6.1 and 6.2
18 | Verify firm has established procedures for identifying training needs; ensure personnel are trained to perform assigned responsibilities | ISO 13485:2016: 6.2 | 21 CFR 820.25(b) | ISO 13485:2016: 6.2 | - Review Procedures: Request and review the organization's procedures for identifying training needs and ensuring personnel are trained to perform their assigned responsibilities. Confirm that these procedures are documented and implemented as required by ISO 13485:2016 section 6.2. - Training Records: Review training records to confirm that they are maintained and updated regularly. Ensure that these records include information on the training provided, the dates of training, and the personnel who received the training. Verify that the records demonstrate compliance with the training requirements as required by ISO 13485:2016 section 6.2(d).
19 | Determine if executive management ensures adequate and effective quality system is implemented. Ensure management is committed to and communicates importance of meeting customer requirements, regulatory requirements, and QMS. | ISO 13485:2016: 5.1(a), 5.2, 5.5.3 |  | ISO 13485:2016: 5.1(a), 5.2, 5.5.3 | - Interview Executive Management: Conduct interviews with executive management to assess their commitment to the QMS. Verify that they ensure the QMS is adequately and effectively implemented and maintained as required by ISO 13485:2016 section 5.1(a). - Subsystem Review: Review the performance of other QMS subsystems, such as design control, document control, production and process controls, and CAPA. Verify that these subsystems are functioning effectively and align with the overall QMS objectives.
Design & Development / Design Controls
1 | Verify products are subject to design controls | ISO 13485:2016: 7.1, 7.3 | 21 CFR 820.30(a) | ISO 13485:2016: 7.1, 7.3 | - Review Design Control Procedures: Request and review the organization's design control procedures to ensure they comply with ISO 13485:2016 sections 7.1 and 7.3
2 | Verify design control and risk management procedures are established and applied | ISO 13485:2016: 7.3 | 21 CFR 820.30(a) - (j) | ISO 13485:2016: 7.3 | - Review Design Control Procedures: Request and review the organization's design control procedures to ensure they comply with ISO 13485:2016 section 7.3. Confirm that these procedures are documented, implemented, and maintained effectively.
3 | Ensure design and development stages are identified; confirm that review, verification, validation, and design transfer activities at each stage are appropriate; verify responsibilities for design and development are defined | ISO 13485:2016: 7.3.2 | 21 CFR 820.30 | ISO 13485:2016: 7.3.2 | - Review Procedures: Request and review the organization's design and development procedures to ensure they comply with ISO 13485:2016 section 7.3.2. Confirm that these procedures are documented, implemented, and maintained effectively. - Design and Development Planning: Verify that the organization has established and maintains design and development plans that identify the different stages of design and development. Ensure these plans describe or reference the design and development activities at each stage as required by ISO 13485:2016 section 7.3.2. - Stage Identification: Confirm that the design and development stages are clearly identified in the planning documents. These stages should include review, verification, validation, and design transfer activities appropriate to each stage. - Review Activities: Verify that the organization conducts formal, documented reviews at appropriate stages of the design and development process. Ensure that these reviews assess the adequacy of design requirements and identify any necessary changes or improvements. - Verification Activities: Confirm that design verification activities are performed to ensure that design outputs meet design input requirements. Review documentation of verification activities to ensure they are appropriate for each stage. - Validation Activities: Ensure that design validation activities are conducted under defined operating conditions to ensure devices conform to user needs and intended uses. Verify that validation activities are appropriate for each stage and are documented as required. - Design Transfer Activities: Confirm that design transfer activities are planned and executed to ensure that the design is correctly translated into production specifications. Review documentation to ensure that design transfer is conducted appropriately at the relevant stage. - Responsibilities for Design and Development: Verify that the responsibilities for design and development activities are clearly defined and assigned. Ensure that the organization has documented the roles and responsibilities of personnel involved in each stage of the design and development process.
5 | Review a project design & development plan, responsibilities, and interfaces | ISO 13485:2016: 7.3.2 | 21 CFR 820.30(b) | ISO 13485:2016: 7.3.2 | - Assess Design and Development Plan: Obtain and assess a specific project’s design and development plan to verify it includes all necessary elements and complies with the standard. Key Elements to Assess in the Plan: Milestones: Verify that the plan includes clearly defined milestones for each stage of the design and development process. Ensure these milestones are appropriate and achievable, providing checkpoints for progress and quality assessment. Phases: Confirm that the plan outlines distinct phases of the design and development process. Each phase should be well-defined with specific objectives, activities, and deliverables. Responsibilities: Ensure that the plan clearly defines the responsibilities of individuals and teams involved in the design and development process. Verify that these responsibilities are assigned and understood by the relevant personnel. Risk Management: Assess the incorporation of risk management activities within the design and development plan. Verify that risk analysis, evaluation, and control measures are identified and implemented at appropriate stages. Interfaces: Confirm that the plan describes the interfaces between different groups or activities involved in the design and development process. Ensure that interactions and communication channels are clearly defined and managed.
6 | Verify design & development plan is updated, reviewed, and approved | ISO 13485:2016: 7.3.2 | 21 CFR 820.30(b) | ISO 13485:2016: 7.3.2 | - Review Plan Revisions: Request and review the organization’s procedures for updating design and development plans to ensure they comply with ISO 13485:2016 section 7.3.2 and 21 CFR 820.30(b). Verify that these procedures are documented, implemented, and maintained effectively. - Review and Approval Procedures: Verify that the organization has established procedures for the review and approval of design and development plan updates. Ensure that these procedures require formal review and approval by designated individuals. - Documented Approval: Confirm that updates to the design and development plan are reviewed and approved according to the established procedures. Ensure that the approval is documented with signatures and dates of the individuals responsible. - Document Review: Review the organization’s design and development plan and its revisions, ensuring they include detailed information on updates, reviews, and approvals as required by ISO 13485:2016 section 7.3.2
7 | Confirm design input requirements were established, reviewed, and approved; ensure customer requirements are captured; ensure inputs include functional, performance, safety, and statutory and regulatory requirements | ISO 13485:2016: 7.2.1, 7.3.3 | 21 CFR 820.30(c) | ISO 13485:2016: 7.2.1, 7.3.3 | - Assess Design Input Requirements: Obtain and assess the specific design input requirements for a project to verify they were established, reviewed, and approved according to the procedures. Key Elements to Assess in Design Inputs: Intended Use: Confirm that the design input requirements address the intended use of the device, ensuring they align with customer expectations and regulatory requirements. Functional, Performance, and Safety Requirements: Verify that the design inputs include detailed functional, performance, and safety requirements. Ensure these requirements are comprehensive and appropriate for the intended use of the device. Statutory and Regulatory Requirements: Ensure that the design inputs incorporate all applicable statutory and regulatory requirements. Confirm that the organization has identified and included relevant standards and regulations. User and Patient Needs: Verify that the design inputs capture the needs of users and patients, ensuring that the device will be safe and effective in its intended use. Other Essential Requirements: Confirm that the design inputs address any other essential requirements specific to the device and its intended use.
8 | Incomplete, ambiguous, and/or conflicting requirements were addressed | ISO 13485:2016: 7.3.3 | 21 CFR 820.30(c) | ISO 13485:2016: 7.3.3 | - Review Procedure: Request and review the organization’s procedures for identifying and addressing incomplete, ambiguous, and conflicting design input requirements to ensure they comply with ISO 13485:2016 section 7.3.3. - Assess Resolutions: Obtain and review records of identified incomplete, ambiguous, or conflicting design input requirements and their resolutions to ensure they were properly addressed according to the procedures. Key Elements to Assess in Resolutions: Identification of Issues: Confirm that the procedures include a mechanism for identifying incomplete, ambiguous, or conflicting design input requirements. Ensure that these issues are documented and tracked. Resolution Process: Verify that the organization has a defined process for resolving incomplete, ambiguous, or conflicting requirements. Ensure that the process includes steps for clarification, validation, and approval of the resolutions. Documentation of Resolutions: Ensure that the resolutions of identified issues are documented, including the rationale for the resolution, the individuals involved, and the approval of the resolution. Review and Approval: Confirm that the resolutions of incomplete, ambiguous, or conflicting requirements are reviewed and approved by designated individuals. Ensure that the approval is documented with signatures and dates.
9 | Confirm design & development outputs are established, verifiable, reviewed, and approved | ISO 13485:2016: 7.3.4(a), (c) | 21 CFR 820.30(d) | ISO 13485:2016: 7.3.4(a), (c) | - Review Procedure: Request and review the organization’s procedures for establishing, verifying, reviewing, and approving design and development outputs to ensure they comply with ISO 13485:2016 sections 7.3.4(a), (c). - Assess Design Outputs: Obtain and assess specific design and development outputs for projects to verify they were established, verified, reviewed, and approved according to the procedures. Key Elements to Assess in Design Outputs: Established Outputs: Confirm that design outputs are established and documented, including drawings, specifications, labeling, packaging, work instructions, and instructions for use (IFUs). Verifiability: Verify that design outputs are verifiable and traceable to design input requirements. Ensure that the outputs meet the specified requirements and are sufficient for the intended use of the device. Review and Approval: Ensure that design outputs are reviewed and approved by designated individuals. Confirm that the review and approval process is documented with signatures and dates.
10 | Ensure design & development outputs are appropriate for purchasing, production, and servicing | ISO 13485:2016: 7.3.4(b) | 21 CFR 820.30(d) | ISO 13485:2016: 7.3.4(b) | - Review Procedure: Request and review the organization’s procedures for ensuring that design and development outputs are appropriate for purchasing, production, and servicing to ensure they comply with ISO 13485:2016 section 7.3.4(b) Key Elements to Assess: Appropriateness for Purchasing: Confirm that design outputs provide sufficient detail for purchasing requirements, including specifications for materials, components, and services. Ensure that purchasing data is clear and meets quality requirements. Suitability for Production: Verify that design outputs include necessary information for production, such as manufacturing instructions, equipment specifications, and process parameters. Ensure that the outputs support consistent production and product quality. Adequacy for Servicing: Ensure that design outputs include servicing requirements, such as maintenance instructions, service procedures, and spare parts lists. Confirm that servicing information is detailed and supports effective device maintenance and repair.
11 | Verify essential design & development outputs are identified | ISO 13485:2016: 7.3.4(d) | 21 CFR 820.30(d) | ISO 13485:2016: 7.3.4(d) | - Review Procedure: Confirm that the design and development procedures document the identification of essential outputs. Verify that these procedures include or reference criteria for identifying essential outputs necessary for the safe and proper use of the medical device. - Review drawings, specifications, labeling, packaging, work instructions, IFUs: Ensure that all essential design outputs, such as drawings, specifications, labeling, packaging, work instructions, and Instructions for Use (IFUs), are properly documented and referenced in the design and development files. Confirm that these outputs meet the input requirements and are in a form suitable for verification and validation.
12 | Confirm acceptance criteria is referenced by design & development outputs and was defined prior to design verification and design validation activities | ISO 13485:2016: 7.3.4(c), 7.3.6 | 21 CFR 820.30(d) & (f) | ISO 13485:2016: 7.3.4(c), 7.3.6 | - Review Procedure: Verify that procedures are in place for defining and documenting acceptance criteria for design outputs. Ensure these criteria are established and documented prior to verification and validation activities. Confirm that procedures specify the requirement to reference acceptance criteria in design outputs. - Review drawings, specifications, labeling, packaging, work instructions, IFUs: Ensure that all design outputs, including drawings, specifications, labeling, packaging, work instructions, and Instructions for Use (IFUs), reference the established acceptance criteria. Confirm that these outputs are adequately detailed to allow verification and validation against the criteria. - Review verification activities: Confirm that design verification activities include checks against the established acceptance criteria. Ensure that the criteria were defined and documented prior to conducting verification activities. Review records of verification activities to verify compliance with documented acceptance criteria.
13 | Determine if design verification confirmed design outputs met design input requirements | ISO 13485:2016: 7.3.6 | 21 CFR 820.30(f) | ISO 13485:2016: 7.3.6 | - Review procedure: Verify that procedures are in place for design verification, ensuring that design outputs are confirmed to meet design input requirements. Confirm that the procedure includes criteria for verifying the alignment between outputs and inputs and the documentation required for verification activities. - Review verification activities: Review records of design verification activities to ensure that they demonstrate that design outputs have been checked and confirmed to meet design input requirements. Verify that verification activities are documented, including the methods used, results obtained, and the conclusion that outputs meet inputs.
14 | Confirm design validation results prove device met predetermined user needs and intended uses | ISO 13485:2016: 7.3.7 | 21 CFR 820.30(g) | ISO 13485:2016: 7.3.7 | - Review procedure: Verify that procedures for design validation are in place and include criteria for ensuring that the device meets predetermined user needs and intended uses. Confirm that the procedure outlines the methods for validation, documentation requirements, and responsibilities for conducting validation activities. - Review validation activities: Examine records of design validation activities to ensure they demonstrate that the device meets user needs and intended uses under defined operating conditions. Verify that validation activities are documented, including the methods used, results obtained, and conclusions that the device meets the specified requirements.
15 | Confirm design validation did not leave unresolved discrepancies | ISO 13485:2016: 7.3.7 | 21 CFR 820.30(g) | ISO 13485:2016: 7.3.7 | - Assess design and specification changes: Verify that design validation activities have been conducted thoroughly, ensuring that all discrepancies identified during validation are resolved before final approval. Confirm that any changes to the design or specifications as a result of validation are documented and reviewed. - Review validation records: Examine validation records to ensure they include documentation of any discrepancies found and the actions taken to resolve them. Verify that there are no unresolved discrepancies that could impact the device's ability to meet user needs and intended uses.
16 | if required by national or regional regulations, confirm clinical evaluations and/or evaluation of device performance were performed | ISO 13485:2016: 7.3.7 | 21 CFR 820.30(g) | ISO 13485:2016: 7.3.7 | - Review procedure: Verify that procedures are in place for conducting clinical evaluations and/or evaluation of device performance as required by national or regional regulations. Confirm that the procedure outlines the criteria, methodology, and documentation requirements for these evaluations. - Review evaluation data: Examine records of clinical evaluations and/or performance evaluations to ensure they have been conducted according to the established procedures and regulatory requirements. Verify that the evaluation data is complete, accurate, and demonstrates that the device meets performance and safety requirements.
17 | If device contains software, confirm software was validated | ISO 13485:2016: 7.3.2, 7.3.7 | 21 CFR 820.30(g), 820.75 | ISO 13485:2016: 7.3.2, 7.3.7 | - Ensure software component has satisfied design, validation, and change control requirements: ISO 13485:2016: 7.3.2: Verify that software design and development plans include appropriate phases for validation and change control. Confirm that the software component is integrated into the overall design and development plan and is subject to the same review, verification, and approval processes as hardware components. ISO 13485:2016: 7.3.7: Review records of software validation activities to ensure that the software has been validated under defined operating conditions. Confirm that validation includes testing the software under actual or simulated use conditions to ensure it meets user needs and intended uses.
18 | Determine if initial production units (or equivalents) were used for design validation | ISO 13485:2016: 7.3.7 | 21 CFR 820.30(g) | ISO 13485:2016: 7.3.7 | - Review procedure: Verify that the organization's design validation procedures specify the use of initial production units, lots, or batches (or their equivalents) for validation activities. Confirm that the procedure outlines the criteria for selecting these units and the validation methods to be used. Ensure the procedure includes requirements for conducting design validation under defined operating conditions and that it includes protocols for simulating actual use conditions. - Evaluate prototype/production records: Review records to confirm that initial production units, lots, or batches (or their equivalents) were used during the design validation process. Ensure that the records document the validation activities performed, the results obtained, and any discrepancies or issues identified during validation.
19 | Confirm risk management activities were performed | ISO 13485:2016: 7.1; ISO 14971:2019 | 21 CFR 820.30(g) | ISO 13485:2016: 7.1; ISO 14971:2019 | - Review procedure: Verify that the organization's procedures for risk management are established, documented, and maintained. Ensure these procedures align with the principles and requirements of ISO 14971:2019. Confirm that the procedure encompasses risk analysis, risk evaluation, risk control, and the evaluation of residual risk acceptability. - Review risk management file: Assess the risk management file to ensure it includes comprehensive documentation of risk management activities, as required by ISO 14971:2019. Verify that the file contains records of risk analysis, risk evaluation, and risk control measures implemented. Ensure that any residual risks are evaluated and deemed acceptable.
20 | Confirm design changes were controlled and validated (or where appropriate, verified) | ISO 13485:2016: 7.3.2, 7.3.6, 7.3.9 | 21 CFR 820.30(i), 820.70(b), 820.75(c) | ISO 13485:2016: 7.3.2, 7.3.6, 7.3.9 | - Review procedure: Verify that the organization has established procedures for the identification, documentation, validation, or verification of design changes before implementation. Ensure that the procedures address the verification of design changes to confirm that design outputs meet design input requirements. Confirm that the procedures include the review, approval, and control of design changes, ensuring that all changes are documented and validated (or verified, where appropriate). - Review design changes and documentation decisions: Assess the documentation for design changes to ensure all changes have been properly identified and documented. Confirm that design changes have undergone appropriate verification activities to ensure compliance with design input requirements. Check that design changes have been reviewed, approved, and controlled according to documented procedures. Ensure validation activities have been performed when necessary.
21 | Confirm design changes have been reviewed for effect on components and product previously made | ISO 13485:2016: 7.3.2, 7.3.6, 7.3.9 | 21 CFR 820.30(i), 820.70(b) | ISO 13485:2016: 7.3.2, 7.3.6, 7.3.9 | - Review design changes and documentation decisions: Verify that the organization has procedures in place to identify, document, and control design changes, ensuring they are reviewed for their potential impact on components and products previously made. Ensure that design changes are verified to meet design input requirements, with an assessment of how these changes might affect existing components and products. Confirm that all design changes have been reviewed, approved, and controlled according to documented procedures, with particular attention to their impact on previously manufactured products and components.
22 | Determine if design reviews were conducted at appropriate stages of design & development | ISO 13485:2016: 7.2.2, 7.3.2, 7.3.5 | 21 CFR 820.30(e) | ISO 13485:2016: 7.2.2, 7.3.2, 7.3.5 | - Review procedure and design review documentation: Verify that procedures are established to conduct design reviews at appropriate stages, ensuring customer requirements are incorporated and addressed. Confirm that design and development planning includes provisions for conducting design reviews at each critical stage of the design process. Ensure design reviews are conducted according to the planned arrangements and documented procedures, including a review of design outputs to ensure they meet design input requirements. Check that design review procedures comply with regulatory requirements, ensuring reviews are planned and conducted at appropriate stages, with documentation of review results.
23 | Confirm design review attendees were appropriate for stage and included independent reviewer | ISO 13485:2016: 7.3.2, 7.3.5 | 21 CFR 820.30(e) | ISO 13485:2016: 7.3.2, 7.3.5 | - Review design review documentation: Verify that the design and development planning includes identifying appropriate design review stages and ensuring suitable participants are involved. Confirm that design reviews include representatives of functions concerned with the design stage being reviewed, as well as an independent reviewer not directly responsible for the design. Check that the design review procedures comply with regulatory requirements by including participants who are independent of the design process and ensuring documentation of review results and attendee lists.
24 | Determine if design was correctly transferred to production | ISO 13485:2016: 7.3.2, 7.3.8 | 21 CFR 820.30(h) | ISO 13485:2016: 7.3.2, 7.3.8 | - Review procedure: Ensure that the design and development planning includes the necessary steps for design transfer, detailing how design outputs are translated into production specifications. Confirm that there are procedures to ensure the design is correctly transferred to production, including verification that production specifications match design outputs. - Review device specifications: Check that device specifications are documented and detailed within the design and development plan. Confirm that the device specifications used in production are consistent with the approved design outputs and meet all required criteria.
25 | Ensure Tech File contains design control documentation | ISO 13485:2016:7.3.10 | 21 CFR 820.30(b) - (j) | ISO 13485:2016: 7.3.10 | - Review Technical File: Verify that the Technical File includes all necessary documentation related to design control activities. This includes records of design and development planning, design inputs, design outputs, design reviews, design verification, design validation, design transfer, design changes, and design history.
Corrective & Preventive Actions (CAPA)
1 | Verify CAPA procedures comply with regulatory requirements | ISO 13485:2016: 4.1, 4.2, 8.5 | 21 CFR 820.100(a) | ISO 13485:2016: 4.1, 4.2, 8.5 | - Review procedures: Verify that the CAPA procedures are documented as part of the QMS, ensuring they align with the organization’s processes and include the necessary controls and responsibilities. Confirm that the CAPA procedures are part of the documented information required by the QMS, including the establishment, implementation, and maintenance of these procedures. Ensure that the procedures address the necessary steps for corrective actions to eliminate the causes of nonconformities and preventive actions to eliminate the causes of potential nonconformities. This includes: Corrective Actions: Investigating the cause of nonconformities, determining the corrective actions needed, and implementing and recording these actions. Preventive Actions: Identifying potential nonconformities, assessing the need for action to prevent occurrence, and implementing and recording preventive actions.
2 | Verify non-conforming product and CAPA procedures determine the need for investigation and notification | ISO 13485:2016: 8.3, 8.5 | 21 CFR 820.90(a), 820.100(a)(2) | ISO 13485:2016: 8.3, 8.5 | - Review procedures: Verify that the non-conforming product procedures ensure the identification, documentation, evaluation, segregation, and disposition of non-conforming products. Ensure that these procedures determine the need for an investigation and notification, including: Evaluating the non-conformance and its impact on the product, process, or system; documenting the decision-making process regarding whether an investigation is necessary, and ensuring the notification of relevant personnel about the non-conforming product and the actions taken.
3 | Verify non-conforming product and CAPA procedures define responsibilities for review and disposition | ISO 13485:2016: 8.3, 8.5; | 21 CFR 820.90(b)(1) | ISO 13485:2016: 8.3, 8.5; 21 CFR 820.90(b)(1) | - Review Nonconformance Procedures: Ensure the procedures address control actions for nonconforming products detected before and after delivery, including rework, re-evaluation, and documentation of these actions. Verify the inclusion of corrective and preventive actions, ensuring the causes of nonconformities are determined, actions are planned and implemented to prevent recurrence, and the effectiveness of these actions is verified. - Review CAPA Procedures: Ensure that corrective and preventive action (CAPA) procedures define responsibilities for the review and disposition of nonconforming products, as required by ISO 13485:2016 clause 8.5 and 21 CFR 820.90(b)(1). (21 CFR 820.90(b)(1): Confirm that the procedures specify the responsibility for review and the authority for the disposition of nonconforming product, and that the disposition process is well-documented. Ensure records include justifications for the use of nonconforming products and the signature of the individual(s) authorizing the use.) - Review Documentation: Ensure records include evaluations of nonconformities, decisions on disposition, rework, or other corrective actions, and verification of the effectiveness of these actions.
4 | Ensure that procedures for rework, retesting, and re- evaluation of nonconforming product exist and are followed | ISO 13485:2016: 8.3, 8.5 | 21 CFR 820.90(b)(2) | ISO 13485:2016: 8.3, 8.5 | - Review Rework Procedures: Ensure that the procedures address actions to be taken when nonconforming products are detected, including rework processes and criteria for accepting reworked products. - Review Records of Nonconforming Products: erify that records include details of rework performed, results of retesting, and decisions made during re-evaluation. Ensure that the records demonstrate compliance with specified requirements and document any corrective actions taken.
5 | Verify that appropriate records of quality problems have been created and used | ISO 13485:2016: 8.3, 8.5 | 21 CFR 820.100(a)(1) | ISO 13485:2016: 8.3, 8.5 21 CFR Part 820.35(a) | Review records of acceptance activities, production test failures, returned products, service records, complaints - Review Records of Acceptance Activities: Confirm that records address acceptance activities for identifying nonconforming products and corrective actions. - Review Production Test Failures: Check documentation for production test failures and corrective actions taken. - Review Returned Product Records: Inspect records of returned products for details of issues found and corrective actions taken. - Review Service Records: Examine service records for documentation of issues and corrective actions. - Review Complaint Records: Ensure records include: The name of the device. The date the complaint was received. Device identifiers. Complainant details. Nature and details of the complaint. Corrective actions taken. Reply to the complainant. - Ensure Documentation of Justification for Not Investigating Similar Complaints: Check that records document the reason for not performing an investigation if a similar complaint has already been investigated. [QMSR 21 CFR 820.35(a)]: Confirm that documentation justifies the decision not to investigate similar complaints and includes the rationale.
6 | Determine if trend analysis data indicates quality problems; determine if data used for CAPA decisions | ISO 13485:2016: 8.1, 8.2.5, 8.4, 8.5 | 21 CFR 820.100(a)(1), 820.250 | ISO 13485:2016: 8.1, 8.2.5, 8.4, 8.5 | - Review Procedures for Trend Analysis: Confirm that procedures define methods for collecting and analyzing data to identify trends that indicate quality problems. - Review Records of Incoming Products and Components: Check records for data on incoming products and components, focusing on trends indicating recurring issues. - Review Testing Records: Inspect testing records for data on product performance, identifying trends that indicate quality problems. - Review Statistical Process Control (SPC) Data: Check SPC records for trends indicating deviations from specified requirements or recurring quality problems. - Assess Use of Trend Analysis Data in CAPA Decisions: Ensure that trend analysis data is systematically used to identify the root causes of quality problems and implement effective corrective actions. - Verify Effectiveness of Trend Analysis Procedures: Assess the effectiveness of trend analysis in identifying and addressing quality issues, ensuring continuous improvement of the quality management system.
7 | Verify CAPA data is complete, accurate, and timely; compare results across multiple data sources to identify quality problems | ISO 13485:2016: 8.4, 8.5 | 21 CFR 820.100(a)(1) | ISO 13485:2016: 8.4, 8.5 | - Compare Results Across Multiple Data Sources: Review data from different sources such as incoming product inspections, in-process tests, final product inspections, customer complaints, and service records. - Use Data Tables to Determine Sampling Plan: Develop and review data tables to determine sampling plans that adequately represent the population of interest, - Assess Effectiveness of CAPA Data Analysis: Evaluate the effectiveness of data analysis methods in identifying root causes and guiding corrective and preventive actions.
8 | Verify appropriate statistical techniques are implemented | ISO 13485:2016: 8.1, 8.2.5, 8.4 | 21 CFR 820.100(a)(1), 820.250 | ISO 13485:2016: 8.1, 8.2.5, 8.4 | - Review Procedures for Statistical Techniques: Confirm that the quality management system includes procedures for using statistical techniques to ensure product quality. - Review Techniques Used: Confirm that statistical techniques are used to monitor and measure the performance of the quality management system. erify that data analysis and evaluation are performed using statistical techniques to identify trends and areas for improvement.
9 | Verify device failure investigations determine root cause | ISO 13485:2016: 8.3, 8.5 | 21 CFR 820.100(a)(2) | ISO 13485:2016: 8.3, 8.5 | - Review Procedures for Device Failure Investigations: Confirm that procedures address the identification and control of nonconforming products, including investigations to determine the root cause. Verify that corrective and preventive actions procedures include steps for investigating device failures to identify root causes. - Review Investigation Records: Verify that records include information on the identification, evaluation, and investigation of nonconforming products. Ensure that records of corrective and preventive actions document the investigation process and root cause determination. - Verify Thoroughness of Investigations: Assess the thoroughness of investigations into nonconforming products, ensuring that root causes are identified and addressed. Verify that investigations into quality issues include comprehensive root cause analysis and lead to effective corrective actions.
10 | Verify failure investigations are commensurate with risks | ISO 13485:2016: 8.3, 8.5 | 21 CFR 820.100(a)(2), 820.90(b) | ISO 13485:2016: 8.3, 8.5 | - Review Procedures: Confirm that procedures outline the criteria for determining when an investigation is required and how its scope should correspond to the level of risk. Verify that procedures include steps for implementing corrective actions based on the severity of the failure and its impact on patient safety. - Review Investigations: Evaluate completed investigations to verify that they were conducted thoroughly, especially for higher-risk failures, and that they identified the root cause effectively. Check that investigations have led to appropriate corrective actions, particularly where the risk was significant, and ensure these actions are documented.
11 | Verify controls exist to prevent non-conforming product from being released | ISO 13485:2016: 8.3 | 21 CFR 820.90(b) | ISO 13485:2016: 8.3 | - Review investigations: Confirm that investigations document the effectiveness of controls in place to prevent the release of non-conforming products and identify any gaps that could lead to non-conforming products being released. - Review Non-Conformance Records: Examine non-conformance records to ensure that they detail the specific controls used to prevent the release of non-conforming products and document their effectiveness.
12 | Verify appropriate actions were taken for quality problems | ISO 13485:2016: 8.2.5, 8.5.2, 8.5.3 | 21 CFR 820.100(a)(3), 820.100(a)(5); 820.100(a)(4), 820.100(b) | ISO 13485:2016: 8.2.5, 8.5.2, 8.5.3 | - Review Procedure: Verify that the procedures outline how to assess the significance of quality problems, determine the root cause, and implement corrective and preventive actions. Ensure the procedures are aligned with the requirements for monitoring and measuring product conformity and managing non-conformities. - Review CAPA Records: Review records to verify that they document the identification of quality problems, the root cause analysis, the actions taken to correct and prevent the issues, and the verification of the effectiveness of these actions. Ensure that records include all necessary information to demonstrate compliance with the CAPA process.
13 | Determine CAPA actions were effective, verified, validated, documented, and implemented appropriately | ISO 13485:2016: 8.5 | 21 CFR 820.100(a)(4), 820.100(a)(5), 820.100(b) | ISO 13485:2016: 8.5 | - Review procedure: Confirm that the procedures specify how CAPA actions are to be validated and verified to ensure they effectively address the identified quality problems and prevent recurrence. Ensure that the procedures require the documentation of these activities. - Review CAPA records: Review records to confirm that they document the verification and validation of CAPA actions, including evidence that the actions taken effectively addressed the root cause of the quality problems. Ensure that records include all necessary information to demonstrate that the CAPA process was followed correctly and that the actions taken were appropriate.
14 | Verify CAPAs and nonconformities were disseminated to personnel responsible for ensuring quality and prevention of problems | ISO 13485:2016: 8.3, 8.5 | 21 CFR 820.100(a)(6) | ISO 13485:2016: 8.3, 8.5 | - Review CAPA and non-conformance records: Confirm that the records show evidence of communication regarding CAPAs and nonconformities to the appropriate personnel. Verify that the communication was timely and included all necessary details for those responsible to take appropriate actions.
15 | Verify quality issues and CAPAs were disseminated for Management Review | ISO 13485:2016: 5.6.3, 8.3, 8.5 | 21 CFR 820.100(a)(6), 820.100(a)(7) | ISO 13485:2016: 5.6.3, 8.3, 8.5 | - Review procedure: Confirm that procedures include a requirement for presenting quality issues and CAPAs during management review meetings. The procedures should specify the frequency and content of reports to management. - Review CAPA records: Check that CAPA records indicate that information was presented to management as part of the review process. Verify that the records show management’s involvement in reviewing CAPAs and their outcomes.
16 | Verify firm has procedures for handling complaints and investigation of advisory notices / recalls; ensure provisions exist to feed in CAPA system | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.2.3 | 21 CFR 820.100, 820.198 | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.2.3 21 CFR Part 820.35(a) | - Review procedures: Verify that the firm's procedures comply with the QMSR, specifically ensuring that complaint handling and investigation activities are documented and incorporated into the CAPA system for continuous improvement. Confirm that procedures detail the process for receiving, documenting, and evaluating complaints, including steps for determining if further investigation is necessary. erify that procedures cover the investigation of advisory notices and recalls, ensuring that root cause analysis is performed and that findings are documented and fed into the CAPA system. Verify that the procedures align with the specific regulatory requirements outlined in 21 CFR 803.17, ensuring that MDR processes meet all applicable federal regulations.
Medical Device Reporting
1 | Verify MDR procedures comply with regulatory requirements | ISO 13485:2016: 8.5.1 | 21 CFR 803.17 | ISO 13485:2016: 8.5.1, 21 CFR 803.17 | - Review procedures: Confirm that the procedures provide clear guidance on how to handle MDRs, including the identification, documentation, and timely reporting of adverse events.
2 | Verify firm maintains MDR event files that comply with regulatory requirements | ISO 13485:2016: 8.5.1 | 21 CFR 803.18 | ISO 13485:2016: 8.5.1, 21 CFR 803.18 | - Review MDR files: Confirm that the MDR files are maintained in a manner that supports the identification, documentation, and management of adverse events, ensuring they align with the requirements of ISO 13485:2016. Verify that the files comply with the specific regulatory requirements outlined in 21 CFR 803.18, including the maintenance of records related to MDR events, ensuring all necessary documentation is preserved and accessible for regulatory review.
3 | Confirm appropriate MDR information is identified, reviewed, reported, documented, and filed | ISO 13485:2016: 8.5.1; | 21 CFR 803, 820.198(d) | ISO 13485:2016: 8.5.1; 21 CFR 803 | - Review MDR files: Verify that the firm has processes in place for identifying, reviewing, and documenting MDR information. Ensure that the documentation supports accurate reporting and timely submission of MDRs. Confirm that MDR information is reported in compliance with 21 CFR 803, including the required time frames and specific reporting criteria.
4 | Ensure firm is effective in identifying MDR reportable events | ISO 13485:2016: 8.5.1; | 21 CFR 803 | ISO 13485:2016: 8.5.1; 21 CFR 803 | - Review procedures: Verify that procedures include clear criteria and processes for identifying events that may be MDR reportable. Ensure that these procedures facilitate the timely and accurate identification of such events. Confirm that the procedures comply with 21 CFR 803, specifying the criteria for what constitutes an MDR reportable event and the steps for identifying such events. - Review MDR files: Check that MDR files document the identification process for MDR events, including how the events were recognized and classified. Verify that MDR files are complete and include all necessary documentation as required by 21 CFR 803, demonstrating the firm’s effectiveness in identifying reportable events. - Review Complaints & Returned Products: Confirm that the firm consistently reviews complaints and returned products for signs of MDR reportable events and that these reviews are documented. Ensure that any potential MDR events identified through complaints and returned products are correctly reported and documented as per 21 CFR 803 requirements.
5 | Ensure firm has established procedures for receiving, reviewing, and evaluating complaints | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1 | 21 CFR 820.198(a) - (c) | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1 21 CFR Part 820.35(a) | - Review procedures: Confirm that the procedures detail the process for receiving and documenting complaints, including initial assessments to determine if further investigation is necessary. verify that all procedures are clearly documented, comprehensive, and fully implemented to ensure they meet FDA’s requirements. Identify any gaps where procedures may be lacking in clarity, completeness, or consistency with the regulatory expectations. Ensure that the complaint handling process effectively captures all necessary information, is consistently applied, and that any issues are appropriately addressed and resolved. Verify that the procedures cover the review and evaluation of complaints, ensuring that root cause analysis is performed where required and that findings are documented.
6 | Verify firm maintains complaint files and that they are reasonably accessible | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1 | 21 CFR 820.198(a), (f), (g) | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1 21 CFR Part 820.35(a) | - Review complaint records: Confirm that the firm has established processes to ensure that complaint files are maintained accurately and include all necessary details. Verify that complaint files contain documented evidence of complaint handling, including investigations and corrective actions, and that they are accessible for review. Ensure that these records are organized, complete, and stored in a manner that makes them reasonably accessible for review and inspection.
7 | Confirm that complaints are evaluated to determine if an event should be a MDR | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.2.3, 8.5.1 | 21 CFR 803, 820.198(a)(3) | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.2.3, 8.5.1 21 CFR Part 820.35(a) | - Review Procedures and Complaint Records: Confirm that procedures outline the steps for evaluating each complaint to assess whether it involves an event that should be reported as an MDR. Verify that the evaluation process includes criteria for determining the seriousness of the event, whether it resulted in harm or posed a significant risk to health. Check that findings from the evaluation are documented, and necessary actions are taken if an MDR report is required. Verify that the procedures outline specific steps for assessing whether a complaint involves a reportable event, that staff are trained to follow these steps, and that all necessary documentation is completed and filed correctly when an MDR determination is made. Ensure that any decision not to report an event as an MDR is properly justified and documented.
8 | Ensure complaint investigations include the device name, date of complaint, device identification number, contact information of complainant, details of complaint, date and results of investigation, any corrective actions, and replies to complainant | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1 | 21 CFR 820.198(e) | ISO 13485:2016: 7.2.3, 8.2.1, 8.2.2, 8.5.1 21 CFR Part 820.35(a) | - Review complaint records: Confirm that complaint records include the name of the device involved, the date the complaint was received, and a unique identification number for the device. Verify that the records contain complete contact information for the complainant, along with detailed descriptions of the complaint. Ensure that the date and results of the investigation are documented, along with any corrective actions taken and the responses provided to the complainant. Verify that the records contain: The device name The date the complaint was received The device identification number or unique device identifier (UDI) Contact information of the complainant Detailed description of the complaint The date and results of the investigation Any corrective actions taken The responses or replies provided to the complainant
Reports of Corrections & Removals (C&R)
1 | Verify C&R procedures comply with regulatory requirements |  | 21 CFR 806 | 21 CFR 806 | - Review procedures: Ensure that the procedures for reporting corrections and removals include clear criteria for determining when a report to the FDA is required. This should encompass scenarios where actions are taken to reduce risk to health or remedy a violation of the act. Check if the procedure outlines a process for submitting reports within the 10-working day timeframe as required by 21 CFR 806.10(b).
2 | Examine records of corrections and/or removals of product |  | 21 CFR 806 | 21 CFR 806 | - Determine if removal was initiated by firm: Assess whether the records clearly indicate that the removal of the product was initiated by the firm. The records should detail the decision-making process and the reasons for initiating the removal, particularly in relation to reducing risk to health or remedying a violation of the act. Ensure that there is a clear linkage between the recorded event and the firm's decision to initiate the removal.
3 | Verify reporting requirements are implemented |  | 21 CFR 806 | 21 CFR 806 | - Review procedures: Ensure that all reporting activities are thoroughly documented. This includes maintaining records of all submitted reports, including the report numbers, dates of submission, and confirmation of receipt by the FDA. Confirm that the firm has implemented its reporting procedures effectively. This includes ensuring that all relevant personnel are trained and that reports are consistently submitted within the required 10-working day timeframe after a correction or removal is initiated.
4 | Identify C&R actions not identified or initiated by firm |  | 21 CFR 806 | 21 CFR 806 | - Identify events not identified by firm: Assess whether the firm has records of corrections and removals that were identified or initiated by external parties, such as customers, healthcare providers, or the FDA, rather than by the firm itself. This includes reviewing any communications or reports from these external sources that prompted a corrective or removal action. Determine whether the firm's monitoring processes are sufficient to identify potential issues that may require a correction or removal. This includes reviewing how the firm monitors adverse events, customer complaints, and other indicators that could signal the need for corrective action.
5 | Confirm existence file of non-reportable corrections and removals |  | 21 CFR 806.20 | 21 CFR 806.20 | - Review non-reportable C&R files: Conduct a thorough review to confirm that the firm maintains complete and accurate files for all non-reportable corrections and removals. These files should include all required details, such as the product's brand name, UDI, event description, corrective or removal actions taken, and justifications for non-reporting. Confirm that the firm retains non-reportable C&R records for the required duration—2 years beyond the expected life of the device, even if the manufacturer or importer has ceased operations. Ensure that these records are securely stored and accessible for review by regulatory authorities if needed.
Medical Device Tracking
1 | Identify all manufactured or imported devices that require tracking |  | 21 CFR 821.20 | 21 CFR 821.20 | - Review product listings: identify all Class II and Class III devices that meet the criteria specified in 21 CFR 821 and therefore require tracking. This includes devices where failure could result in serious adverse health consequences, devices intended to be implanted for more than one year, or life-supporting/life-sustaining devices used outside of a device user facility. Confirm that the firm has received and documented any tracking orders issued by the FDA in response to premarket notification submissions or premarket approval applications. This documentation should be readily accessible and tied to the relevant products in the firm’s records.
2 | Verify tracking procedures comply with regulatory requirements |  | 21 CFR 821.25(c) | 21 CFR 821.25(c) | - Review procedures: Conduct a thorough review of the firm's written standard operating procedures (SOPs) for tracking devices. These procedures must outline the processes for collecting, maintaining, and auditing the data required by the FDA for device tracking. Ensure that the procedures cover all necessary aspects, including data collection, record maintenance, system modifications, and quality assurance.
3 | Verify firm performs internal audits of tracking system per timeframes specified in regulations |  | 21 CFR 821.25(c)(3) | 21 CFR 821.25(c)(3) | - Review procedures: Examine the firm’s written procedures to ensure they include a clear schedule and methodology for conducting internal audits of the tracking system. These procedures should specify that audits are to be conducted at least every six months for the first three years of device distribution and at least annually thereafter. Review the records of completed audits to confirm that the firm has conducted internal audits of its tracking system according to the specified timeframes—every six months for the first three years and annually thereafter. Ensure that each audit was performed on time and that no audits were missed or delayed. Confirm that all audit findings are documented and that any identified issues are followed up with appropriate corrective actions. Review records to ensure that the firm has implemented changes or improvements based on audit findings and that these actions are documented and tracked.
Production & Process Controls (main subsystem) (P&PC)
1 | Verify product realization processes are planned; confirm that risk management occurs throughout product realization | ISO 13485:2016: 7.1; ISO 14971:2019 | 21 CFR 820.70 | ISO 13485:2016: 7.1; ISO 14971:2019 | - Review procedures: Confirm that the procedures for product realization include detailed planning, addressing all aspects of the process, from design and development to production and delivery, ensuring that the processes meet specified requirements. Verify that risk management is an integral part of the product realization process, with procedures in place to identify, assess, and mitigate risks at each stage, ensuring that the safety and performance of the product are maintained.
2 | Verify planning of product realization is consistent with requirements of other processes of QMS | ISO 13485:2016: 7.1 | 21 CFR 820.30, 820.50, 820.80, 820.181 | ISO 13485:2016: 7.1 | - Review Procedures: Confirm that the procedures ensure the planning of product realization is conducted in a manner consistent with the requirements of the QMS. The planning should consider all stages of the product lifecycle and ensure that the requirements of design, purchasing, production, and documentation processes are integrated. - Review Product Realization Documents: Examine product realization documents to confirm that they reflect the planned activities and are consistent with QMS processes, including risk management, design controls, supplier management, and production controls.
3 | Verify requirements have been defined for suppliers, contractors, and consultants; ensure suppliers, contractors, and consultants are selected on ability to meet requirements | ISO 13485:2016: 7.1, 7.4.2 | 21 CFR 820.50(a) | ISO 13485:2016: 7.1, 7.4.2 | - Review Procedures: Confirm that the procedures include requirements for planning and managing supplier, contractor, and consultant relationships as part of the overall product realization process. Ensure that these relationships are integrated into the QMS and that responsibilities are clearly defined. Verify that procedures specify the criteria for selecting suppliers, contractors, and consultants based on their ability to meet product and service requirements. This includes the evaluation, selection, monitoring, and re-evaluation of these entities to ensure ongoing compliance. - Review Supplier Records: Examine supplier records to verify that they document the planning, evaluation, and management of supplier relationships, ensuring alignment with QMS processes. Review records to confirm that suppliers, contractors, and consultants have been selected based on their ability to meet defined requirements. Check that evaluations, audits, and monitoring activities are documented and that any issues are addressed and resolved.
4 | Ensure firm maintains records of acceptable suppliers, contractors, and consultants | ISO 13485:2016: 7.4.1 | 21 CFR 820.50(a)(3) | ISO 13485:2016: 7.4.1 | - Review Supplier Records: Verify that records are maintained for all suppliers, contractors, and consultants that meet the firm’s acceptance criteria. Confirm that these records include evaluations, agreements, and any relevant performance monitoring documentation
5 | Verify that data supporting supplier requirements is maintained; verify that suppliers, contractors, and consultants agree to notify firm of changes in products and/or services | ISO 13485:2016: 7.4 | 21 CFR 820.40, 820.50(a)(3), (b) | ISO 13485:2016: 7.4 | - Review records: Confirm that records demonstrate the firm’s evaluation and selection process for suppliers, contractors, and consultants. Ensure that these records include detailed requirements and agreements specifying the need for notification of any changes in products or services.
6 | Verify procedures for identifying product during all stages of receipt, production, distribution, and installation are in place | ISO 13485:2016: 7.5.8, 7.5.9 | 21 CFR 820.60 | ISO 13485:2016: 7.5.8, 7.5.9 21 CFR Part 820.35(c) | - Review procedures: Confirm that procedures include detailed methods for identifying products, ensuring traceability and control throughout the entire lifecycle, from receipt of materials to final installation. Ensure that the procedures address labeling, identification numbers, and any other relevant identification methods. Check that the procedures are clearly defined, regularly reviewed, and effectively implemented by the staff, with particular attention to ensuring that identification methods (e.g., labeling, barcoding, UDI) are accurate.
7 | Ensure firm maintains procedures and records for traceability of each unit, lot, or batch of finished devices and components (may not be required for all devices) | ISO 13485:2016: 7.5.9 | 21 CFR 820.65 | ISO 13485:2016: 7.5.9 21 CFR Part 820.35(c) | - Review procedures: Confirm that procedures include detailed methods for maintaining traceability of each unit, lot, or batch, ensuring that records accurately reflect the production history of finished devices and components. Verify that the traceability system is capable of identifying the relationship between the finished product and its components or materials. Review the firm's procedures and a sample of production records to ensure that the firm maintains adequate procedures and records for the traceability of each unit, lot, or batch of finished devices and their components. Verify that these procedures are properly documented, clearly outline the traceability requirements, and ensure that each device and component can be accurately traced through all stages of production and distribution. - Review production records: Review production records to confirm that they include traceability information, such as control numbers, batch numbers, and the history of each finished device and its components.
8 | Select a process to review |  |  |  | Selection criteria: - CAPA indicators of process issues - process for higher risk device - degree of risk for process to cause device failures - lack of familiarity and experience with process - process used for multiple devices - variety in process technologies - processes not covered during previous inspections
9 | Verify process is controlled and monitored | ISO 13485:2016: 7.5, 7.6, 8.2.5, 8.2.6, 8.4 | 21 CFR 820.50, 820.70(a), 802.70(e), 820.70(f)- (h), 820.72, 820.75(b), 820.80 | ISO 13485:2016: 7.5, 7.6, 8.2.5, 8.2.6, 8.4 | - Review Specific Procedures and Instructions: Confirm that the process is defined in procedures and work instructions, with clear controls to ensure that it operates within specified limits. Verify that process parameters, equipment settings, and environmental conditions are controlled and monitored to maintain process consistency. - Evaluate Process Monitoring and Control: Verify that monitoring activities, such as in-process inspections, process control charts, and process validation, are in place and functioning effectively. Check that monitoring results are analyzed to detect trends, and that corrective actions are taken when process performance does not meet specified requirements. - Review In-Process and Finished Device Acceptance Activities: Review acceptance activities to confirm that they are based on established criteria and that these criteria are aligned with product specifications and regulatory requirements.
10 | Verify the equipment used has been adjusted, calibrated, and maintained | ISO 13485:2016: 7.5 | 21 CFR 820.70(g)(3), 820.72(a), 820.70(g)(1) | ISO 13485:2016: 7.5 | - Review Equipment Records: Confirm that records of equipment adjustments, calibrations, and maintenance are complete and up-to-date. Verify that the equipment used in the production process meets the required specifications and is capable of producing valid results. - Review Procedures: Verify that procedures are established for adjusting, calibrating, and maintaining equipment. Ensure that these procedures specify the frequency and methods for performing these activities and that they are adhered to by trained personnel.
11 | Identify control and oversight activities; ensure control of inspection, measuring , test equipment, and calibration | ISO 13485:2016: 7.6, 8.4 | 21 CFR 820.50(a)(2), 820.72 | ISO 13485:2016: 7.6, 8.4 | - Review production, equipment, maintenance, and calibration records: Verify that procedures and records for controlling inspection, measuring, and test equipment are established and maintained. Ensure that calibration activities are performed at appropriate intervals and that the equipment is capable of providing accurate and reliable measurements. Confirm that data related to control and oversight activities is analyzed and used to monitor the effectiveness of the quality management system. Ensure that this data supports continual improvement and corrective actions where necessary. In-Process Acceptance Criteria & Acceptance: Ensure that in-process inspection and acceptance criteria are well-defined and adhered to, with proper oversight and control of the equipment used. Finished Device Acceptance Criteria & Acceptance: Verify that finished device acceptance activities are controlled and that the equipment used is calibrated and maintained according to the specified requirements. Environmental Control Systems: Review records related to environmental control systems to ensure they are functioning correctly and are subject to appropriate oversight and maintenance. Contamination Control Systems: Check that contamination control systems are in place and that equipment related to these systems is calibrated, maintained, and controlled to prevent any adverse effects on product quality.
12 | Verify firm has established procedures for production and process changes; ensure changes are verified or validated, as needed | ISO 13485:2016: 7.3.9, 7.5.6 | 21 CFR 820.70(b), 820.75(c) | ISO 13485:2016: 7.3.9, 7.5.6 | - Review Procedures: Confirm that procedures address the requirements for design and development changes, ensuring that these changes are reviewed, verified, and validated to maintain the integrity of the product. Verify that the procedures include controls for production and service provision changes, ensuring that these changes are managed in a controlled manner and are subject to appropriate verification or validation.
13 | Review production records to identify rejects and/or non-conformances | ISO 13485:2016: 8.3 | 21 CFR 820.70 | ISO 13485:2016: 8.3 | - Review Production Records: Confirm that production records include details of non-conformances, including the nature of the non-conformance, the quantity affected, and the actions taken. Verify that the records demonstrate compliance with the procedures established for controlling non-conforming products.
14 | Verify that defects, rejects, non-conformances, and removal of materials were handled properly | ISO 13485:2016: 8.3 | 21 CFR 820.50, 820.70(h), 820.90, 820.100 | ISO 13485:2016: 8.3 | - Review Material and Production Records: Confirm that records detail how defects, rejects, and non-conformances were identified and addressed. Verify that any removal of materials is documented, including the reason for removal and the actions taken to prevent recurrence.
15 | Ensure processes that cannot be fully verified are validated | ISO 13485:2016: 7.5.6 | 21 CFR 820.75(a) | ISO 13485:2016: 7.5.6 | - Review procedures: Confirm that procedures clearly define the processes that require validation and outline the steps for conducting and documenting validation activities. Identify Processes That Cannot Be Verified: Ensure that all such processes are identified in the firm's documentation and that there is a clear rationale for why these processes require validation rather than verification. Review Validation Records: • Ensure that the validation records meet all necessary requirements and demonstrate compliance with regulatory standards. • Verify that all operators involved in the validated processes have documented qualifications and training relevant to the processes they perform. • Ensure that any changes to validated processes are controlled and documented, including revalidation where necessary. • Confirm that all instruments used in the validated processes are calibrated and maintained according to documented schedules and requirements. • Verify that equipment used in validated processes is properly installed, adjusted, and maintained to ensure consistent performance. • Check that product specifications are clearly defined before the validation process begins and that these specifications are consistently met during validation. • Ensure that test sampling and validation plans are based on statistically valid methods and that the sampling is adequate to ensure process control. • Confirm that process tolerance limits are adequately challenged during validation to ensure that processes remain within acceptable limits under all anticipated conditions.
16 | Ensure automated or software driven processes are validated for intended uses | ISO 13485:2016: 7.5.6 | 21 CFR 820.70(i) | ISO 13485:2016: 7.5.6 | Review Validation Records: Confirm that the validation records demonstrate that automated or software-driven processes are validated to ensure they perform consistently and as intended. • Check that the software validation includes documentation showing that all software-driven processes have been tested under expected operating conditions and that the software performs as intended. • Ensure that any changes to the software or automated processes are controlled, and that revalidation is performed as necessary to ensure continued compliance and effectiveness. • Confirm that validation activities are fully documented, including the results of all validation tests, any identified issues, and how they were addressed to ensure the software meets its intended use.
17 | Verify that validations are documented and conducted by qualified personnel | ISO 13485:2016: 7.5.6 | 21 CFR 820.75(b)(1) | ISO 13485:2016: 7.5.6 | - Review Procedures and Validation Records: Confirm that procedures require validations to be carried out by personnel with documented qualifications. Ensure that validation activities are appropriately recorded and include evidence of personnel qualifications. Check that all validation activities are fully documented, including details of the personnel involved, the methods used, and the results obtained. Ensure that the records demonstrate that personnel conducting validations have received the necessary training and possess the qualifications required for the validation activities they perform. Ensure that procedures outline the requirements for documenting validations and the qualifications needed for personnel conducting validations.
18 | Review personnel records to document personnel are trained per manufacturing processes and aware of potential defects | ISO 13485:2016: 6.2 | 21 CFR 820.20(b)(2), 820.25, 820.70, 820.70(d), 820.75(b)(1) | ISO 13485:2016: 6.2 | - Review Personnel Records: Confirm that personnel records show evidence of training in the manufacturing processes they are involved in. Verify that training includes awareness of potential defects and their impact on product quality.
19 | Ensure that monitoring and control methods, data, date performed, individuals performing the process, and the major equipment used is documented | ISO 13485:2016: 7.1, 8. | 21 CFR 820.75(b)(2) | ISO 13485:2016: 7.1, 8. | - Review Validation Records: Confirm that the records include detailed documentation of the monitoring and control methods used during the process, aligning with the planning of product realization. Verify that all relevant data, including test results, and records of monitoring activities are documented accurately.
20 | Determine linkages to other processes | ISO 13485:2016: 4.1, 4.2 | 21 CFR 820.20, 820.25, 820.30, 820.40, 820.72, 820.90, 820.100, 820.180 | ISO 13485:2016: 4.1, 4.2 | - Review Procedures: Confirm that the QMS processes are interconnected and integrated, ensuring that linkages between processes are clearly defined to maintain the effectiveness of the QMS. Verify that the documentation and control of records ensure traceability and alignment across different processes within the QMS.
21 | Ensure the infrastructure and work environment are appropriate and controlled | ISO 13485:2016: 6.3, 6.4 | 21 CFR 820.70(c), (f), (g) | ISO 13485:2016: 6.3, 6.4 | - Review Procedures: Confirm that the firm’s procedures detail the necessary infrastructure, including buildings, workspace, associated utilities, process equipment, and supporting services, ensuring they are suitable for achieving conformity to product requirements: Verify that procedures address the work environment, including necessary cleanliness, hygiene, and conditions required to ensure product quality. This includes environmental controls, contamination controls, and any other relevant conditions.
22 | Confirm that maintenance schedules, routine inspections, and adjustments to equipment occur | ISO 13485:2016: 6.3, 7.5.1, 7.5.6, 7.6 | 21 CFR 820.70(g) | ISO 13485:2016: 6.3, 7.5.1, 7.5.6, 7.6 21 CFR Part 820.35(c), 21 CFR Part 820.45 | - Review Procedures: Verify that procedures include requirements for maintaining the infrastructure, ensuring that equipment and supporting services are suitable for achieving product conformity. Confirm that procedures define the necessary production and process controls, including equipment maintenance and inspection requirements to ensure consistent product quality. Ensure that procedures require validation of processes where verification is not feasible, including controls for equipment used in such processes. Verify that procedures cover the calibration and maintenance of inspection, measuring, and test equipment, ensuring that it is fit for its intended purpose. Verify that these procedures ensure all equipment is regularly maintained, inspected, and adjusted as needed to maintain its proper function. Check that maintenance records are up-to-date, accessible, and demonstrate compliance with the established schedules and procedures, ensuring that any deviations or corrective actions are properly documented and addressed.
23 | Verify procedures are in place for contamination control and cleanliness | ISO 13485:2016: 6.4.2, 7.5.2 | 21 CFR 820.70(e) | ISO 13485:2016: 6.4.2, 7.5.2 | Review Procedures: Confirm that procedures include requirements for controlling contamination during production, especially where cleanroom environments or controlled conditions are necessary to meet product specifications. Verify that procedures address the cleanliness of the product and production environment, specifying controls to prevent contamination that could affect the product’s quality or safety.
24 | Determine if verification of purchased products is adequate | ISO 13485:2016: 7.4.3, 8.4 | 21 CFR 820.50(a)(2), 820.80(b) | ISO 13485:2016: 7.4.3, 8.4 | Review Procedures: Confirm that the firm has established procedures for verification of purchased products to ensure they meet specified requirements. This includes defining the criteria for acceptance and the methods for verifying compliance. Verify that the procedures include monitoring and measuring activities to ensure purchased products continue to meet quality requirements during production.
25 | Ensure procedures define receiving, in-process, and final acceptance activities. | ISO 13485:2016: 7.5.11, 8.4 | 21 CFR 820.80(a) - (d) | ISO 13485:2016: 7.5.11, 8.4 | Review Procedures: Confirm that the procedures outline the specific acceptance activities for each stage, including the criteria for receiving, in-process, and final acceptance. The procedures should detail the methods for verifying that products meet specified requirements before moving to the next stage. Ensure that the procedures include provisions for monitoring and measuring activities during the receiving, in-process, and final acceptance stages to verify product conformity and maintain quality throughout production.
26 | Confirm receiving, in-process, and final acceptance activity records exist | ISO 13485:2016: 8.4 | 21 CFR 820.80(e) | ISO 13485:2016: 8.4 | Review Records: Ensure that records are maintained for all acceptance activities at each stage of production. These records should include details of the acceptance criteria, results of inspections and tests, and any nonconformances identified and addressed.
27 | Verify that procedures exist and that acceptance status of product is indicated | ISO 13485:2016: 7.1, 8.2.6 | 21 CFR 820.86 | ISO 13485:2016: 7.1, 8.2.6 | Review Acceptance Criteria: Ensure that the planning of product realization includes the establishment of acceptance criteria that align with customer and regulatory requirements. Verify that procedures for monitoring and measurement of products include clear criteria for acceptance or rejection at various stages of production. Review Procedures: Check that procedures include methods for clearly marking or labeling products to indicate their acceptance status, preventing mixups and ensuring that only conforming products move to the next stage of production or distribution. Review Product Identification: Review product labeling, tags, or other identification methods to ensure that the acceptance status is clearly visible and consistent with the firm’s procedures. Ensure that the acceptance status is traceable through the firm’s records, linking the status to specific inspection and test results, as well as any rework or corrective actions taken. Verify that the acceptance status is consistently applied and communicated across all relevant departments, including production, quality control, and distribution.
28 | Ensure procedures define labeling activities, including integrity, inspection, storage, operations, and control numbers | ISO 13485:2016: 7.5.11 | 21 CFR 820.120 | ISO 13485:2016: 7.5.11 | Review Procedures: Verify that the procedures include detailed steps for ensuring the integrity of labeling throughout the production process. This includes ensuring that labels are correctly printed, applied, and inspected for accuracy before use.
29 | Confirm that product packaging and shipping containers adequately protect device during processing, storage, handling, shipping, and distribution | ISO 13485:2016: 7.5.11 | 21 CFR 820.130 | ISO 13485:2016: 7.5.11 | Review Procedures: Verify that procedures cover the selection, design, and testing of packaging and shipping containers to ensure they meet the requirements for protecting the device during processing, storage, handling, shipping, and distribution.
30 | Verify procedures exist to prevent mix-ups, damage, deterioration, contamination, or other adverse effects to product during handling | ISO 13485:2016: 7.5.11 | 21 CFR 820.140, 820.150 | ISO 13485:2016: 7.5.11 | Review Procedures: Verify that procedures cover all aspects of product handling, including the prevention of mix-ups, damage, deterioration, and contamination. Confirm that these procedures are adequately documented and implemented throughout the organization.
31 | Verify procedures exist for product distribution; confirm distribution records include name and address of consignee, identification and quantity shipped, date of shipment, and identification numbers | ISO 13485:2016: 4.2.3, 7.1, 7.5.8, 7.5.9.2, 7.5.11 | 21 CFR 820.160 | ISO 13485:2016: 4.2.3, 7.1, 7.5.8, 7.5.9.2, 7.5.11 21 CFR Part 820.35(c) | Review procedures and distribution records: review the firm's procedures and a sample of distribution records to verify that procedures for product distribution are in place and compliant with 21 CFR Part 820.35(c). Verify that the firm maintains documented procedures for the control of records related to product distribution, ensuring that they are readily retrievable and protected against damage or loss. Confirm that distribution procedures cover all aspects of the distribution process, including the identification and tracking of product, to ensure that only products meeting specified criteria are distributed. Procedures should also ensure that distribution records include the consignee’s name and address, product identification, quantity shipped, date of shipment, and relevant identification numbers.
32 | Ensure installation and inspection procedures exist (if applicable); verify installation records are maintained | ISO 13485:2016: 7.5.3 | 21 CFR 820.170 | ISO 13485:2016: 7.5.3 | Review Procedures: Verify that procedures are in place for the installation of devices, including steps for ensuring that installation is performed correctly and that the device functions as intended after installation. Procedures should also include inspection criteria to confirm proper installation. Check a sample of installation records to ensure that they include all required details, such as the date of installation, the individual performing the installation, the device identification, and the results of any inspections performed.
33 | Ensure servicing procedures exist (if applicable); verify servicing records are maintained | ISO 13485:2016: 7.5.4; | 21 CFR 820.200 | ISO 13485:2016: 7.5.4; 21 CFR Part 820.35(b) | Review procedures and servicing records: Verify that procedures are in place for the servicing of devices, including steps for performing and documenting all servicing activities. Procedures should include criteria for ensuring that servicing is conducted appropriately and that the device remains safe and effective after servicing. Review a sample of servicing records to verify that these records are accurately maintained and include all required information. Specifically, confirm that servicing records document: The nature of the service performed The date of service The identification of the serviced device (e.g., serial number, UDI) The details of the service performed, including any replacement parts or adjustments made The results of any inspections or tests conducted after servicing
34 | Verify firm identifies, verifies, protects, and safeguards customer property under its care | ISO 13485:2016: 7.5.10 |  | ISO 13485:2016: 7.5.10 | Review Procedures: Confirm that procedures are in place for the proper identification and verification of customer property when it is received by the firm. Procedures should include steps for documenting the condition of the property, as well as measures for protecting and safeguarding it from damage, loss, or misuse. Check a sample of records to ensure that they include all required details, such as the identification of the customer property, the verification of its condition, and the protective measures applied.
Sterilization Process Controls
1 | Review sterilization process procedures; verify sterilization process is validated | ISO 13485:2016: 7.5.7 | 21 CFR 820.75(a), (c) | ISO 13485:2016: 7.5.7 | Review Procedures: Confirm that procedures outline the sterilization process in detail, including parameters such as time, temperature, and pressure. Procedures should specify the methods for validating the sterilization process to ensure it consistently achieves the desired Sterility Assurance Level (SAL).
2 | Review sterilization control and monitoring activities; ensure processes, equipment, and calibration are current | ISO 13485:2016: 7.5.5 | 21 CFR 820.50, 820.70(a), (c),(e), (f), (g), (h), 820.72, 820.75(b), 820.80 | ISO 13485:2016: 7.5.5 | Review Sterilization Records: Verify that records demonstrate ongoing control and monitoring of the sterilization process. Ensure that the records include details of the sterilization cycles, including time, temperature, pressure, and other relevant parameters.
4 | Review production records for sterilization failures; ensure integration with CAPA system | ISO 13485:2016: 7.5.5, 4.2.5 | 21 CFR 820.75(b) | ISO 13485:2016: 7.5.5, 4.2.5 | Review Sterilization Records: Ensure that production records for sterilization include detailed documentation of any failures, including the nature of the failure, the time, date, and specific batch or lot affected. Verify that these records provide a clear and accurate account of the sterilization process. Confirm that the documentation of sterilization failures is maintained as part of the overall quality management system and that these records are accessible and available for review. Ensure that the records are complete and accurately reflect the production environment.
5 | Ensure sterilization failures were handled properly | ISO 13485:2016: 7.5.5, 4.2.5 | 21 CFR 820.70(g)(3), 820.72(a), 820.70(g)(1) | ISO 13485:2016: 7.5.5, 4.2.5 | Review Sterilization Records: Ensure that sterilization records include detailed documentation of any failures, specifying the nature of the failure, affected batches or lots, and the date and time of the occurrence. Verify that the documentation reflects a clear understanding of the sterilization process and the issues encountered. Confirm that records of sterilization failures are maintained as part of the overall quality management system, ensuring that they are accessible and properly archived. Verify that these records demonstrate compliance with the firm’s procedures for handling sterilization failures. Confirm that the firm’s response to sterilization failures meets the requirements outlined in ISO 13485:2016 clause 7.5.5, including the documentation and resolution of issues to maintain product quality.
6 | Review personnel records to document personnel are qualified and trained with implemented sterilization activities | ISO 13485:2016: 6.2, 4.2.5 | 21 CFR 820.25, 820.70(d), 820.75(b) | ISO 13485:2016: 6.2, 4.2.5 | Review Personnel Records: Verify that personnel records demonstrate appropriate education, background, and training for individuals performing sterilization activities. Confirm that the records include details of the training provided, its relevance to the sterilization processes, and any certifications or qualifications held by the personnel. Ensure that personnel records are maintained as part of the quality management system and are accessible for review. Confirm that these records are kept up to date and include all relevant documentation related to training and qualifications for sterilization activities. Confirm that the training records show a clear link between the training provided and the specific sterilization activities performed. Ensure that training programs are reviewed regularly and updated as needed to address changes in processes or regulations.
7 | Ensure automated or software driven sterilization processes are controlled and validated | ISO 13485:2016: 7.5.5, 7.5.7, 4.2.5 | 21 CFR 820.70(i) | ISO 13485:2016: 7.5.5, 7.5.7, 4.2.5 | Review Validation Records: Verify that the validation records demonstrate that sterilization processes are adequately controlled to prevent contamination or other adverse effects on the product. Ensure that all steps of the automated or software-driven process are documented, monitored, and controlled. Confirm that the sterilization process validation records show that the process has been fully validated to meet predetermined specifications. This includes ensuring that the sterilization process consistently achieves the required Sterility Assurance Level (SAL) without compromising product performance or packaging integrity. Ensure that records of validation activities are maintained within the quality management system and are accessible for review. These records should include detailed documentation of the validation process, results, and any subsequent control measures implemented for the automated or software-driven sterilization processes. Verification of Process Control and Validation: Confirm that the records reflect ongoing control and monitoring of the sterilization process, including any software updates or changes to automation. Verify that the validation process includes testing under simulated or actual operating conditions to ensure consistent and reliable sterilization.
Purchasing Controls
1 | Review supplier evaluation procedures | ISO 13485:2016: 7.4.1 | 21 CFR 820.50 | ISO 13485:2016: 7.4.1 | Review Procedures: Confirm that the procedures define the criteria and methods used for the initial selection, evaluation, and ongoing monitoring of suppliers, contractors, and consultants. This includes assessing their ability to provide products or services that consistently meet specified quality requirements.
2 | Ensure suppliers are evaluated for ability to meet specified requirements | ISO 13485:2016: 7.4.1 | 21 CFR 820.50(a)(1) | ISO 13485:2016: 7.4.1 | Review Procedures: Confirm that the procedures include detailed criteria for assessing the capability of suppliers to provide products or services that meet the firm’s quality requirements. This should include initial qualification and ongoing evaluation processes.
3 | Ensure adequacy of specifications of materials and/or services provided by supplier is confirmed | ISO 13485:2016: 7.4.2 | 21 CFR 820.50(b) | ISO 13485:2016: 7.4.2 | Review Procedures: Confirm that procedures include detailed methods for assessing whether the specifications of materials and/or services provided by suppliers are adequate to meet the firm's quality requirements. Ensure that validation or verification steps are documented.
4 | Confirm purchasing information identifies requirements for approval of product, procedures, processes, and equipment, requirements for personnel qualification, and QMS requirements | ISO 13485:2016: 7.4.2 | 21 CFR 820.50 | ISO 13485:2016: 7.4.2 | Review Purchasing Records: Check that purchasing records detail the necessary approvals for products, procedures, processes, and equipment. Ensure that these records also specify the qualifications required for personnel involved in these processes, as well as the QMS requirements that suppliers must meet. Review Procedures: Confirm that procedures specify the requirements for approval of products, processes, and equipment, as well as personnel qualifications. Ensure that these procedures align with the QMS requirements and are consistently applied.
5 | Verify supplier evaluation records are maintained | ISO 13485:2016: 7.4.1 | 21 CFR 820.50(a)(3) | ISO 13485:2016: 7.4.1 | Review Procedures: Confirm that the procedures outline the process for evaluating suppliers, including the criteria for evaluation and the documentation requirements for maintaining supplier evaluation records. Review Supplier Evaluation Records: Check that the supplier evaluation records include all necessary documentation, such as the evaluation criteria, the results of the evaluation, and any actions taken based on the evaluation. Ensure that these records are up-to-date and stored in accordance with the firm’s procedures.
6 | Determine that verification and acceptance of purchased materials and/or services is adequate | ISO 13485:2016: 7.4.3 | 21 CFR 820.50(a)(2), 820.80(a), 820.80(b) | ISO 13485:2016: 7.4.3 | Review Procedures: Confirm that procedures specify the verification activities that must be conducted for purchased materials and services, including criteria for acceptance and rejection. Review Acceptance Records: Check that acceptance records include all necessary documentation, such as inspection results, testing data, and verification activities, and that they demonstrate conformity with the specified requirements.
Documentation & Records
1 | Review procedures for identification, storage, protection, retrieval, retention time, control, approval, distribution, disposition, and changes of documents and records | ISO 13485:2016: 4.2.4, 4.2.5 | 21 CFR 820.40, 820.180 | ISO 13485:2016: 4.2.4, 4.2.5 21 CFR Part 820.35, 21 CFR Part 820.45 | Review procedures: Confirm that the procedures cover all aspects of document control, including identification, storage, protection, retrieval, and retention time, ensuring that documents remain accessible and legible throughout their retention period. Verify that procedures include processes for approving, distributing, and making changes to documents and records, ensuring that only current, approved versions are in use and that obsolete documents are promptly removed from points of use.
2 | Ensure documents and changes are approved prior to use | ISO 13485:2016: 4.2.4 | 21 CFR 820.40 | ISO 13485:2016: 4.2.4 | Review Procedures: Confirm that the procedures specify that documents and any subsequent changes are reviewed and approved by designated personnel with the appropriate authority, ensuring that only approved documents are used within the QMS. Review Documents and Records: Review a sample of documents and change management records to verify that they have been approved by authorized personnel prior to their use. Confirm that this approval is documented, including the date and signature of the approver. Review Change Management Records: Evaluate change management records to confirm that all changes have undergone a documented review and approval process before being implemented, ensuring traceability and control of document versions.
3 | Verify documents and records are legible and identifiable | ISO 13485:2016: 4.2.4(e), 4.2.5 |  | ISO 13485:2016: 4.2.4(e), 4.2.5 21 CFR Part 820.35, 21 CFR Part 820.45 | Review Documents and Records: Verify that documents and records are clear and readable, ensuring that important details such as text, dates, and signatures are easily discernible. Confirm that documents are properly identified and organized, allowing for quick retrieval and review. The review should confirm that: Text and information on the documents are clear and not obscured by damage, fading, or poor print quality. Each document is appropriately titled and identified, with unique identifiers where necessary (e.g., document numbers, revision levels). The content of the records clearly corresponds to the intended process, product, or activity, with no ambiguity in interpretation.
4 | Ensure documents of external origin are identified with controlled distribution | ISO 13485:2016: 4.2.4(f) |  | ISO 13485:2016: 4.2.4(f) | Review External Documents and Records: Verify that the firm has procedures in place to identify and control the distribution of external documents, such as standards, regulations, customer specifications, and supplier documentation. Check that these documents are clearly marked and tracked to ensure they are distributed to the appropriate personnel and are up to date.
5 | Verify firm maintains a quality system record (QSR) which includes or refers to location of procedures | ISO 13485:2016: 4.2.1(c), (e) | 21 CFR 820.20, 820.40, 820.186 | ISO 13485:2016: 4.2.1(c), (e) | Review Procedures: Verify that the QSR is comprehensive, up-to-date, and properly maintained, including references to all applicable procedures as required by the quality management system.
6 | Confirm that documents and records are retained for required length of time (this includes retention of obsolete controlled documents and records) | ISO 13485:2016: 4.2.1, 4.2.4, 4.2.5 | 21 CFR 820.100(b), 820.180(b), 820.181, 820.184, 820.186, 820.198(a), 820.200(d) | ISO 13485:2016: 4.2.1, 4.2.4, 4.2.5 21 CFR Part 820.35, 21 CFR Part 820.45 | Review Documents and Records: Review records to confirm that they are kept for the required duration and that obsolete documents are managed appropriately. Review Procedures: Verify that the procedures specify the retention periods for documents and records, ensuring that all required information is retained as per the quality management system's requirements. Verify that the procedures clearly define the retention periods for all documents and records, including obsolete controlled documents and records. The review should ensure that: Retention times are explicitly stated in the procedures and align with regulatory requirements. Both active and obsolete documents and records are stored securely and are retained for the entire required period. There is a system in place for tracking retention periods and ensuring timely disposal of documents and records that are no longer required to be kept. Obsolete controlled documents are appropriately marked and stored in a manner that prevents unauthorized use.
7 | Ensure change records are reviewed and approved by the same functions that performed original review and approval | ISO 13485:2016: 4.2.4, 7.3.9 | 21 CFR 820.40(b) | ISO 13485:2016: 4.2.4, 7.3.9 | Review Procedures: Confirm that the procedures clearly outline the requirement for the same functions to be involved in both the original and subsequent change approvals. Review Change Records: Review change records to verify that they have been reviewed and approved by the same functions that performed the original review and approval, ensuring compliance with the quality management system’s requirements.
8 | Verify change records include a description of change, identification of affected documents, approval signatures, approval date, and effective date | ISO 13485:2016: 7.3.9 | 21 CFR 820.40(b) | ISO 13485:2016: 7.3.9 | Review Procedures: Confirm that procedures outline the necessary details to be included in change records, ensuring comprehensive documentation and traceability of changes. Review Change Records: Review change records to verify that they contain a full description of the change, clearly identify the affected documents, and include approval signatures, the date of approval, and the effective date of the change.
9 | Ensure documents are available at point of use and obsolete document are not in use | ISO 13485:2016: 4.2.4(d), (h) | 21 CFR 820.40(a) | ISO 13485:2016: 4.2.4(d), (h) | Review Procedures: Confirm that procedures outline the process for ensuring that only current documents are available at the point of use, and that obsolete documents are properly identified and removed from all points of use. Review Document Distribution: Check that the document distribution process ensures that current documents are accessible where needed and that obsolete documents are not available for use. Review Change Management Records: Review change management records to verify that when documents are updated, the old versions are removed from circulation and the new versions are distributed appropriately.
10 | Verify that firm maintains device specifications for each type of device | ISO 13485:2016: 4.2.1 | 21 CFR 820.181 | ISO 13485:2016: 4.2.1 | Review Production Records: Confirm that production records include detailed specifications for each device, covering aspects such as design, materials, performance requirements, and any other relevant specifications.
11 | Ensure that design and development files contain or make reference to device specifications, production process specifications, quality assurance procedures and specifications (including acceptance criteria), packaging and labeling specifications (including acceptance criteria), and installation, maintenance, and servicing procedures | ISO 13485:2016: 4.2.1 | 21 CFR 820.181(a) - (e) | ISO 13485:2016: 4.2.1 | Review Production Records: Ensure that production records include or reference complete device specifications, production process specifications, and quality assurance procedures, including acceptance criteria for both product quality and packaging/labeling.
12 | Verify that production records are maintained and devices are manufactured according to device specifications; ensure realization processes and product meet requirements | ISO 13485:2016: 7.1, 8.2.6 | 21 CFR 820.184 | ISO 13485:2016: 7.1, 8.2.6 | Review Production Records: Confirm that the planning and implementation of product realization processes are consistent with the device specifications and are effectively documented in the production records. Verify that the production records demonstrate that the products have been subjected to appropriate monitoring and measurement activities to ensure conformity with the acceptance criteria.
13 | Confirm that production records contain or make reference to dates of manufacture, quantity manufactured, quantity released for distribution, acceptance records demonstrating the device was manufactured per device specifications, primary identification label and labeling used for each unit, and device identification and/or control numbers used. | ISO 13485:2016: 8.2.6 | 21 CFR 820.184(a) - (f) | ISO 13485:2016: 8.2.6 | Review Production Records: Confirm that production records accurately reflect the monitoring and measurement activities performed, ensuring that devices meet the specified requirements.
14 | Ensure firm maintains records for education, training, skills, and experience of resources | ISO 13485:2016: 6.2 (e) |  | ISO 13485:2016: 6.2 (e) | Review Training Records: Confirm that the firm maintains up-to-date records of the education, training, skills, and experience of all personnel. These records should include: Documented evidence of relevant education and training. Records of completed training programs and certifications. Details of skills and experience pertinent to assigned responsibilities. Evidence that personnel are qualified to perform their tasks based on their documented qualifications.
15 | Verify firm maintains purchasing and supplier records | ISO 13485:2016: 7.4.1, 7.4.3 | 21 CFR 820.50 | ISO 13485:2016: 7.4.1, 7.4.3 | Review Purchasing and Supplier Records: Confirm that the firm maintains records of: Supplier evaluations and re-evaluations. Purchasing data, including specifications and agreements with suppliers. Documentation of verification and acceptance of purchased products and services. Records demonstrating that suppliers have been selected based on their ability to meet specified requirements. Evidence that purchasing information clearly defines requirements for the product, process, and services, and that this information has been communicated to suppliers.
16 | Ensure sterilization process parameters and records are maintained for each batch; ensure sterilization validation records are maintained | ISO 13485:2016: 7.5.5, 7.5.7 |  | ISO 13485:2016: 7.5.5, 7.5.7 | Review Sterilization Records: Confirm that records include: Specific sterilization process parameters (e.g., time, temperature, pressure). Batch-specific data that documents the adherence to these parameters during each sterilization cycle. Documentation of any deviations or adjustments made during the process, along with justifications and corrective actions taken. Verify that sterilization validation records include: Evidence that the sterilization process has been validated according to established protocols. Records demonstrating that validation activities, including revalidation when necessary, are performed by qualified personnel. Documentation showing that the sterilization process consistently achieves the required sterility assurance level (SAL) and that product quality and packaging integrity are not adversely affected by the process.
Customer Requirements
1 | Review product requirements to ensure that intended use, customer requirements, and regulatory requirements are addressed | ISO 13485:2016: 7.2.2 |  | ISO 13485:2016: 7.2.2 | Review Procedures: Verify that procedures include: Steps for identifying and documenting the intended use of the product. Processes for capturing and analyzing customer requirements to ensure they are fully understood and incorporated into the product design and development stages. Methods for identifying and integrating relevant regulatory requirements, ensuring the product complies with all applicable laws and regulations.
2 | Confirm incoming contracts and orders are reviewed to resolve conflicting information and that customer requirements can be met | ISO 13485:2016: 7.2.2 |  | ISO 13485:2016: 7.2.2 | Review Procedures: Verify that procedures include: Processes for reviewing all incoming contracts and orders to identify any conflicting information. Steps to resolve conflicts in contracts or orders to ensure consistency with customer requirements. Methods for confirming that all customer requirements, as stated in contracts or orders, can be fulfilled before acceptance. Review Incoming Inspection Records: Check that incoming inspection records: Document the review of each contract or order, including the identification of any conflicts and the steps taken to resolve them. Demonstrate that customer requirements were verified and confirmed as achievable before the contract or order was accepted. Ensure that any issues identified during the review process are appropriately documented and resolved in accordance with the established procedures.
3 | Verify that procedures and systems exist for customer communications and feedback; ensure integration with CAPA system | ISO 13485:2016: 7.2.3, 8.2.1 |  | ISO 13485:2016: 7.2.3, 8.2.1 | Review Procedures: Verify that procedures cover: Processes for receiving, documenting, and responding to customer communications, including complaints and feedback. Methods for ensuring that all customer communications are properly logged and tracked through to resolution. Steps for escalating significant customer issues into the CAPA system for corrective and preventive actions. Integration of customer feedback into the CAPA system to ensure that potential quality issues identified through customer communications are addressed systematically. Processes for analyzing customer feedback to identify trends and recurring issues, and for initiating CAPA activities as necessary. Documentation requirements for customer communications and the corresponding CAPA actions taken.
Technical Files
1 | Review technical file procedures | ISO 13485:2016: 4.2.1(d) |  | ISO 13485:2016: 4.2.1(d) | Review Procedures: Verify that procedures include: The structure and content requirements for technical files, ensuring they meet all necessary regulatory and quality standards. Processes for maintaining and updating technical files, ensuring that they remain current and reflect any changes in design, production processes, or regulatory requirements. Methods for controlling access to technical files, ensuring that they are only accessible to authorized personnel. Documentation of all changes made to the technical files, with proper approval and version control to track modifications.
2 | Review documents need to ensure planning, operation, and control of technical file processes | ISO 13485:2016: 4.2.1(d) |  | ISO 13485:2016: 4.2.1(d) | Review Procedures: Verify that procedures include: The structure and content requirements for technical files—Ensure that technical files are structured in a way that meets all necessary regulatory and quality standards, including detailed specifications, risk assessments, design documentation, and compliance with applicable regulations. Processes for maintaining and updating technical files—Confirm that there are procedures in place for regularly updating technical files, ensuring that they remain current and reflect any changes in design, production processes, or regulatory requirements. This includes keeping records of updates and ensuring that the latest versions of documents are readily available. Methods for controlling access to technical files—Verify that access to technical files is restricted to authorized personnel only, with clear procedures for how access is granted and monitored. Ensure that there are safeguards in place to protect the integrity and confidentiality of the technical files. Documentation of all changes made to the technical files—Ensure that any changes made to technical files are properly documented, with clear approval processes and version control. This includes tracking modifications, approvals, and maintaining an audit trail to demonstrate compliance with regulatory and quality requirements.
3 | Select documentation to review |  |  |  | Single product focus: Choose documentation that pertains to a specific product, particularly one that is critical or high-risk. Risk-based: Prioritize documentation related to products or processes that have a higher risk of failure or non-compliance. Result of complaints: Select documentation for review based on products that have generated complaints, ensuring that the issues raised are adequately addressed. Recent project: Focus on documentation from a recently completed project to ensure all regulatory and quality requirements were met. Covers product range: Ensure the selected documentation represents a variety of products within the firm’s portfolio to assess consistency in quality and compliance.
4 | Verify documentation addresses a general description of product, intended use(s), and any variants, accessories, or other devices used in combination with product | ISO 13485:2016: 7.1, 7.2, 7.3.4 |  | ISO 13485:2016: 7.1, 7.2, 7.3.4 | Review technical file records: Ensure the documentation includes a comprehensive description of the product, covering all essential aspects such as design, materials, and functional specifications. Confirm that the documentation clearly outlines the intended use(s) of the product, including any specific conditions or limitations. This should align with the product’s labeling and marketing claims. Review the documentation to ensure it includes details about any variants of the product, accessories, or other devices intended to be used in combination with the product. This ensures that all components are properly integrated and tested together.
5 | Ensure design specifications, standards applied, and results of risk analysis are present | ISO 13485:2016: 7.1, 7.2, 7.3.4 |  | ISO 13485:2016: 7.1, 7.2, 7.3.4 | Review technical file records: Examine the technical file records to ensure that they contain all necessary design specifications, clearly document the standards applied, and include comprehensive results from the risk analysis. Verify that these elements are aligned with the product's intended use and regulatory requirements, and that they support the overall safety and effectiveness of the product.
6 | Confirm that principal requirements have been fulfilled | ISO 13485:2016: 7.1, 7.2, 7.3.4 |  | ISO 13485:2016: 7.1, 7.2, 7.3.4 | - Review technical file records: Examine the technical file records to ensure that they include documentation confirming the fulfillment of all principal requirements related to the product’s design, intended use, risk management, and compliance with applicable standards and regulations. Ensure that these records provide clear evidence that the product meets the necessary safety, performance, and quality standards.
7 | Review techniques used to verify design and validate product(s) clinical data | ISO 13485:2016: 7.1, 7.2, 7.3.4 |  | ISO 13485:2016: 7.1, 7.2, 7.3.4 | - Review technical file records: Examine the technical file records to ensure that they include comprehensive documentation of the design verification and validation techniques, particularly those involving clinical data. Ensure that these records provide evidence of methodical and appropriate use of verification and validation processes, confirming the product's compliance with its intended use and safety standards.
8 | Ensure documentation defines sterilization method and validation | ISO 13485:2016: 7.1, 7.2, 7.3.4 |  | ISO 13485:2016: 7.1, 7.2, 7.3.4 | - Review technical file records: Examine the technical file records to ensure they explicitly define the sterilization method used and include detailed validation records. Verify that these records show the sterilization process has been rigorously tested and validated to ensure it meets safety and regulatory requirements.
9 | Ensure documentation includes instruction manual(s) and labeling | ISO 13485:2016: 7.1, 7.2, 7.3.4 |  | ISO 13485:2016: 7.1, 7.2, 7.3.4 | - Review technical file records: Examine the technical file to ensure it contains all necessary instruction manuals and labeling. Verify that these documents provide clear, concise, and accurate information for the proper use, maintenance, and safety of the product, and that they comply with all relevant standards and regulations.
10 | Verify major subcontractors have been documented | ISO 13485:2016: 7.1, 7.2, 7.3.4 |  | ISO 13485:2016: 7.1, 7.2, 7.3.4 | - Review technical file records: Examine the technical file to ensure it contains comprehensive documentation of all major subcontractors. Verify that the records include detailed information about the subcontractors' qualifications, their specific contributions to the product, and any agreements or contracts in place that outline their responsibilities and obligations.

## Read Me
About this QMSR gap analysis worksheet This worksheet is designed to help organizations conduct a gap analysis to harmonize their quality management systems with FDA's Quality System Regulation (QSR), the Quality Management System Regulation (QMSR), and ISO 13485:2016. The tool facilitates a systematic review of various elements within your quality management system to identify non-conformances (NC), opportunities for improvement (OFI), and other key audit findings. About the QMSR: The QMSR aims to streamline regulations and reduce the burden on device manufacturers. Although similar to ISO 13485, the QMSR includes specific modifications for compliance. It directly incorporates ISO 13485:2016 by reference and ISO 9000:2015 terms and definitions for clarity and consistency. Device manufacturers must update documents and align with QMSR requirements, particularly in risk management and electronic records compliance. The transition period extends to February 2, 2026, with FDA inspections maintaining authority. Worksheet Structure: - Assessment Detail: Detailed descriptions of what needs to be assessed, including specific auditor actions, document reviews, and employee interviews. - Auditor Actions: Specific steps auditors should take to verify compliance with QMSR/ISO requirements. - Observation: Space to record observations made during the audit or assessment. - Evidence: Documentation or evidence that supports the observations. - Outcome (NC, OFI, PP, A): The result of the assessment, categorized as: NC: Non-Conformance, OFI: Opportunity for Improvement, PP: Positive Practice, A: Adequate/Acceptable How to Use: - Ensure you have all relevant documents, such as the Quality Manual, training records, audit reports, and management review minutes. - Familiarize yourself with the specific requirements of FDA QSR, QMSR, and ISO 13485:2016. - Use the "Item" and "Assessment Detail" sections to guide your review of each component of the quality management system. - Follow the "Auditor Actions" to ensure a thorough assessment. - Record your findings in the "Observation" and "Evidence" columns. - Categorize the outcome of each assessment using the NC, OFI, PP, or A categories. - After completing the gap analysis, review the findings to identify areas that require corrective actions, improvements, or further monitoring. - Use the results to update your quality management system to align with regulatory and standard requirements. Contact us for assessment and QMS support. Other Recommended Steps: - Thoroughly review the final QMSR ruling to understand the specific requirements and changes from the previous QSR. Develop and implement a comprehensive training program to ensure your team is well-versed in the new QMSR standards, particularly focusing on the integration of ISO 13485:2016 requirements. - Create a detailed Quality Plan outlining the steps needed to transition your QMS and operation to comply with QMSR. This plan should include actionable steps, assigned responsibilities, deadlines, and progress metrics. Proactively communicate with your critical suppliers about the upcoming changes. Ensure they understand the new requirements and the impact on their operations. Collaborate with them to align their quality systems with the QMSR. Begin updating your documentation and operations to reflect QMSR requirements, focusing on areas such as risk management, electronic records, and compliance procedures. How The FDA Group Can Help: Having helped many firms with QSR and ISO 13485:2016 compliance, we offer a comprehensive suite of QMSR compliance services tailored to support medical device companies in navigating and adhering to the new regulatory requirements. Our services are designed to address the specific challenges posed by the QMSR, leveraging our expertise to facilitate a smooth transition. We help you allocate your resources more efficiently, ensuring that your transition to QMSR compliance is completed within deadlines and budget, without compromising day-to-day operations. Our consultants offer an objective assessment of your quality management system, identifying potential issues and opportunities for improvement that internal teams may overlook. Each QMSR support engagement is tailored precisely to your needs and can include, but is not limited to: - Regulatory Gap Analysis — We conduct an in-depth analysis of your current quality management system against the QMSR, identifying gaps and areas for improvement. - Customized Compliance Strategy — Based on the gap analysis, we develop a tailored strategy to address identified gaps, ensuring your systems are fully compliant with the QMSR requirements. - Risk Management Integration — If desired, our experts guide you in embedding risk management processes throughout your quality management system, ensuring compliance with the QMSR's emphasis on risk-based decision-making. - Documentation Updates — We assist in reviewing and updating your documentation, including SOPs, quality manuals, and records to meet the stringent documentation requirements of the QMSR and ISO 13485:2016. - Employee Training and Development — We provide targeted training programs to ensure your team is well-versed in the QMSR changes, their roles, and responsibilities within the updated system. - Preparation for FDA Inspections — With anticipated changes in the FDA's inspection approach under the QMSR, we prepare you for inspections through mock audits and readiness assessments, ensuring your team is confident and prepared. Contact us for assessment and QMS support. Additional Resources: - Quality Management System Regulation: Final Rule Amending the Quality System Regulation – Frequently Asked Questions - The FDA’s Quality Management System Regulation (QMSR): An Introduction and Compliance Guide Acknowledgments and References: - Quality Management System Regulation: Final Rule Amending the Quality System Regulation – Frequently Asked Questions - QMSR Explained: What FDA QSR & ISO 13485 Harmonization Means for Medical Device Companies - Harmonizing ISO 13485 and FDA QSR 21 CFR 820